## Slide 1
Selected Topics In EMI Test Instrumentation
EMC Compliance
Huntsville, AL
When EMI test equipment was built from tubes, and EMC engineers were built of iron…
Courtesy of Ed Price, retired EMC engineer
Courtesy of Ed Price, retired EMC engineer
Portable EMI receiver advertised in Stoddart Electro Systems Catalog, circa 1967

## Slide 2
Ken Javor
30+ YEARS IN EMI/EMC
Consultant to Government and Industry Industry representative to MIL-STD-461 and MIL-STD-464 Tri-Service Working Groups… ken.javor@emccompliance.com (256) 650-5261

## Slide 3
Measurements of Interest
Current (audio & radio frequency) and time domain
Potential (audio & radio frequency) and time domain
Electric Field, radio frequency
Magnetic Field, audio frequency

## Slide 4
Frequency Domain & Time Domain Measurements
then and now

## Slide 5
Selected Topics in Frequency Domain Measurements
(HP Application Note 150)
The electric field of a point charge, a group of charges; the similarity (inverse square law action-at-distance) of electric and Newtonian gravitational fields

## Slide 6
Selected Topics in Frequency Domain Measurements
The electric field of a point charge, a group of charges; the similarity (inverse square law action-at-distance) of electric and Newtonian gravitational fields
Traditional swept frequency spectrum analyzer block diagram from HP App Note 150
From: Hoffmann, Sidhom, Braun and Russer. A Broadband, Low-Noise Time-Domain System for EMI Measurements through Ka-Band up to 40 GHz (IN Compliance magazine July 2013)
FFT vs. Traditional EMI Receiver

## Slide 7
Does wideband (7 MHz) FFT slice return same results as traditional receiver when signal is broadband and the antenna factor changes rapidly within the 7 MHz slice?
FFT vs. Traditional EMI Receiver,
Comparison Using Biconical Antenna

## Slide 8: Measurement
Source: CS115 impulse generator (30 ns on-time, 2 nanosecond rise/fall-time, rep rate 30 pps driving short stub (BW to 160 MHz)
Receiver: various biconical antennas connected to either an HP 8566B, or the R&S ESU. Receivers sweep 20 - 30 and 30 - 80 MHz.
Only difference between measurements is which receiver is connected to the biconical

## Slide 9: Electro-Metrics EM-6912 Measurement Set-Up
CS115 source driving stub antenna through 6 dB pad
Electro-Metrics EM-6912 biconical connects to either HP8566B or R&S ESU in control chamber

## Slide 10: Test Data - EM-6912
Conclusion:
Excellent correlation!
(above 8566 noise floor)
Both machines use same
  30 - 80 MHz frequency range & scale
  100 dBuV reference level
  10 dB rf attenuation
  vertical scale (10 dB/div)
  100 kHz RBW, 300 kHz VBW
  sweep times set by machines
Raw data: no antenna or other factor correction applied

## Slide 11: Test Data - EM-6912
Conclusion:
Excellent correlation!
(above 8566 noise floor)
Both machines use same
  20 - 30 MHz frequency range & scale
  100 dBuV reference level
  10 dB rf attenuation
  vertical scale (10 dB/div)
  100 kHz RBW, 300 kHz VBW
  sweep times set by machines
Raw data: no antenna or other factor correction applied

## Slide 12: EMCO 3109 Measurement Set-Up
CS115 source driving stub antenna through 6 dB pad
EMCO 3109 biconical connects to either HP8566B or R&S ESU in control chamber

## Slide 13: Test Data - EMCO 3109
Conclusion:
Excellent correlation!
(above 8566 noise floor)
Raw data: no antenna or other factor correction applied
Both machines use same
  30 - 80 MHz frequency range & scale
  100 dBuV reference level
  10 dB rf attenuation
  vertical scale (10 dB/div)
  100 kHz RBW, 300 kHz VBW
  sweep times set by machines

## Slide 14: Test Data - EMCO 3109
Conclusion:
Excellent correlation!
(above 8566 noise floor)
Raw data: no antenna or other factor correction applied
Both machines use same
  20 - 30 MHz frequency range & scale
  100 dBuV reference level
  10 dB rf attenuation
  vertical scale (10 dB/div)
100 kHz RBW, 300 kHz VBW
sweep times set by machines

## Slide 15: MIL-STD-461A Biconical Measurement Set-Up
CS115 source driving stub antenna through 6 dB pad
MIL-STD-461A biconical connects to either HP8566B or R&S ESU in control chamber

## Slide 16: Test Data - MIL-STD-461A Biconical
Conclusion:
Excellent correlation!
(above 8566 noise floor)
Raw data: no antenna or other factor correction applied
Both machines use same
  30 - 80 MHz frequency range & scale
  100 dBuV reference level
  10 dB rf attenuation
  vertical scale (10 dB/div)
  100 kHz RBW, 300 kHz VBW
  sweep times set by machines
ambient

## Slide 17: Test Data - MIL-STD-461A Biconical
Conclusion:
Excellent correlation!
(above 8566 noise floor)
Raw data: no antenna or other factor correction applied
Both machines use same
  20 - 30 MHz frequency range & scale
  100 dBuV reference level
  10 dB rf attenuation
  vertical scale (10 dB/div)
  100 kHz RBW, 300 kHz VBW
  sweep times set by machines
ambient

## Slide 18
Selected Topics in Frequency Domain Measurements
(HP Application Note 150)
The electric field of a point charge, a group of charges; the similarity (inverse square law action-at-distance) of electric and Newtonian gravitational fields

## Slide 19
Selected Topics in Frequency Domain Measurements
(HP Application Note 150)
The electric field of a point charge, a group of charges; the similarity (inverse square law action-at-distance) of electric and Newtonian gravitational fields

## Slide 20
Selected Topics in Frequency Domain Measurements
(HP Application Note 150)
The electric field of a point charge, a group of charges; the similarity (inverse square law action-at-distance) of electric and Newtonian gravitational fields

## Slide 21
John McCloskey
NASA/GSFC Chief EMC Engineer
Code 565
301-286-5498
John.C.McCloskey@nasa.gov
Crest Factor for Thermal Noise
Dec. 18, 2013
Selected Topics in Frequency Domain Measurements
(HP Application Note 150)

## Slide 22: Thermal Noise Floor Limit
k = Boltzmann’s constant = 1.38 x 10-23 J/K
T = temperature in Kelvin
B = bandwidth
Theoretical limit for thermal power noise:
Thermal voltage noise across resistance R:
For T = 298 K and R = 50 Ω (typical for RF receiver inputs):
For B = 1 kHz: VN = 14 nV = -37 dBµV
For B = 10 kHz: VN = 45 nV = -27 dBµV
For B = 100 kHz: VN = 142 nV = -17 dBµV
For B = 1 MHz: VN = 450 nV = -7 dBµV
Theoretical noise floor limits of receiver front end
Real receivers/amplifiers have noise floor above this theoretical limit by a factor called the Noise Figure (NF), typically ~ a few dB for a good low noise amplifier (LNA).
NOTE: These are RMS values. The instantaneous peak values are different.
-174 dBm/Hz
-67 dBuV/Hz

## Slide 23: Peak Detector
Peak detector specified by MIL-STD-461F:
How do we relate peak to rms values for noise floor measurements?
Enter the Crest Factor…

## Slide 24: Crest Factor for Thermal Noise
From Henry Ott’s “Electromagnetic Compatibility Engineering”, Section 8.2, pages 332-334:
Crest factor is the ratio of a waveform’s peak to rms value. For a pure sine wave, it is √2.

## Slide 25: Crest Factor for Thermal Noise (cont.)
Thermal noise follows Gaussian (normal) probability distribution:
μ = mean value = 0
σ = standard deviation unit = 1

## Slide 26: Crest Factor for Thermal Noise (cont.)
|VN| < 3.5*Vrms
99.9% of the time
Maximum |VN| = ~3.5*Vrms
20*log10(3.5) = ~11 dB

## Slide 27: Comparison of Measured Average and Peak Levels
Peak levels consistently ~11 dB higher than AVG levels

## Slide 28: Excerpt from R&S ESIB Datasheet
“Crest factor”