## Slide 1
Directional Couplers

## Slide 2
Directional Couplers
Directional couplers measure forward and/or reverse power in a transmission line. A fraction of the forward and/or reverse power is sampled at the eponymous ports. Thus they are handy for monitoring large amounts of power using low power circuits (CS114, RS103). Alternatively, they can be used to attenuate high power signals without the attendant signal-to-noise ratio degradation inherent in a resistive pad -useful for generating a high SNR low-level field intensity using a high power amplifier, e.g., generating 2 V/m from a 200 Watt amplifier.
They are also useful for gauging the degree of match/mismatch in transmission lines and their loads. Comparing reverse power to forward power yields return loss, which can be used to calculate vswr.

## Slide 3
Directional Couplers
Directional couplers can gauge the degree of match/mismatch in transmission lines and their loads. Comparing reverse power to forward power yields return loss, which can be used to calculate vswr.

## Slide 4
Directional Couplers - Calibration
Forward power / through power (-50 dB)

## Slide 5
Directional Couplers - Calibration
Forward power / through power (-50 dB)
-50 dB is expected value with 50 Ohm load

## Slide 6
Directional Couplers - Calibration
Reverse power / through power (-50 dB) with 50 Ohm load on output

## Slide 7
Directional Couplers - Calibration
Reverse power / through power (-50 dB)
fwd pwr port rd’g
20 dB return loss
20 dB return loss
40 dB return loss

## Slide 8
Directional Couplers - Calibration
Reverse power / forward power (both -50 dB) with output loaded by 50 Ohms;
Reverse power is theoretically non-existent and the ratio should be very small (-many tens of dBs)

## Slide 9
Directional Couplers - Calibration
Reverse power / forward power (both -50 dB)
Reflected power is within 20 dB of forward power; best performance is -40 dB.

## Slide 10
Directional Couplers - Calibration
Reverse power / forward power (both -50 dB) with output open-circuited:
100% reflection means reverse power should equal forward power; no return loss

## Slide 11
Directional Couplers - Calibration
Reverse power / forward power (both -50 dB) with output open-circuited
expect 0 dB

## Slide 12
Calibrating a Directional Coupler:
Using a Spectrum Analyzer and Signal Source in lieu of a network analyzer

## Slide 13
Calibrating a Directional Coupler: Using a Spectrum Analyzer and Signal Source in lieu of a network analyzer
Using NA source set to 0 dBm sweeping 200 – 500 MHz

## Slide 14
Calibrating a Directional Coupler: Using a Spectrum Analyzer and Signal Source in lieu of a network analyzer
Using NA source set to 0 dBm sweeping 200 – 500 MHz;
Output attenuated 50 dB to be commensurate with coupled ports

## Slide 15
Calibrating a Directional Coupler: Using a Spectrum Analyzer and Signal Source in lieu of a network analyzer
Using NA source set to 0 dBm sweeping 200 – 500 MHz;
Lighter trace is fwd pwr port; ideally superimposed on original straight through trace: excellent correlation

## Slide 16
Calibrating a Directional Coupler: Using a Spectrum Analyzer and Signal Source in lieu of a network analyzer
Using NA source set to 0 dBm sweeping 200 – 500 MHz;
Note output open circuited. Lighter trace is rev pwr port; ideally superimposed on fwd pwr trace: correlation within a dB

## Slide 17
Calibrating a Directional Coupler: Using a Spectrum Analyzer and Signal Source in lieu of a network analyzer
-10 dBm from NA source
+15 dBm from NA source

## Slide 18
Calibrating a Directional Coupler: Using a Spectrum Analyzer and Signal Source in lieu of a network analyzer
Fwd pwr -10 dBm from NA source
Rev pwr +15 dBm from NA source
Reverse power should be infinitesimal. In fact it is 2 dB above fwd pwr, but there is a 25 dB difference between fwd and reverse power, so that in fact reverse power is 23 dB below fwd power. That sets up the maximum dynamic range on measuring a good match.

## Slide 19
Measuring VSWR – checking an antenna balun

## Slide 20
Measuring VSWR – checking an antenna balun
10 kHz – 250 MHz directional coupler

## Slide 21
Measuring VSWR – checking an antenna balun
10 kHz – 250 MHz directional coupler

## Slide 22
Measuring VSWR – checking an antenna balun
This set-up compares forward and reverse power to assess vswr.

## Slide 23
Measuring VSWR – checking an antenna balun
Important to keep track of details…

## Slide 24
Measuring VSWR – checking an antenna balun
return pwr
fwd pwr

## Slide 25
Measuring VSWR – checking an antenna balun