## Slide 1: Understanding Radiation Thermometry – Part I
Timothy K. Risch
NASA Armstrong Flight Research Center
November 27, 2015
Narrated by Nicholas DeStefano
Notes:                 Welcome to this NASA NESC Academy lesson on radiative thermometry. In this lesson, we will learn about the theory and practice of radiative thermometry with the primary purpose of determining the temperature of heated bodies using electronic detectors. The course also will cover the historical development of the field, the fundamental physics of radiative surfaces, along with modern measurement methods. Many documented examples to illustrate the concepts are presented throughout this lesson. This is a two-part lesson. Once you have completed this first part, you can proceed to Part II.

## Slide 2: Lesson Plan
Summary of Part 1
Nomenclature
Introduction
History of Radiation Thermometry
Fundamental Physics
Practical Radiation Thermometers
Practical Measurement Techniques
Calibration
Final Thoughts
For Further Reading
Part I
Part II
Notes: Our lesson plan is divided into ten sections, five in Part I and five in Part II. In Part I here, we will introduce a consistent nomenclature used in our lesson followed by an introduction and then history of radiative thermometry. We will then review the fundamental physics behind radiative thermometry. In Part II we will discuss practical radiation thermometers and measurement techniques. Finally we will discuss calibration techniques and then end with a summary of the lesson.

## Slide 3: Preliminaries
An Excel® worksheet containing solutions to all examples are available from Tim Risch at Timothy.K.Risch@NASA.gov
Excel®, Visual Basic for Applications®, and Microsoft® are registered trademarks of Microsoft Corporation.
Notes: All the examples used in this lesson are contained in an Excel® file developed by the author. This file utilizes a library of radiative functions written in Visual Basic for Applications® (VBA) that is also available. To obtain the Excel® file and library, please contact Tim Risch at Timothy.K.Risch@NASA.gov.

## Slide 4: Lesson Plan
Summary of Part 1
Nomenclature
Introduction
History of Radiation Thermometry
Fundamental Physics
Practical Radiation Thermometers
Practical Measurement Techniques
Calibration
Final Thoughts
For Further Reading
Part I
Part II
Notes: Now we will begin Part I.

## Slide 5: Nomenclature - I
speed of light, 2.99792458 × 108 m/s
Planck’s first constant, = 1.191043 × 108 W-m4/m2-sr
Planck’s second constant, = 14,387.75 m-K
Constant in Wien’s displacement law
Constant in equation for maximum blackbody intensity
detector spectral response function for detector I
fraction of spectral intensity between two wavelength- temperature values and
Planck’s constant, 6.626068 × 10-34 J-s
Boltzmann constant, 1.3806503 × 10-23 J/K
spectral emissive intensity of a perfect blackbody at wavelength  and temperature , W/m2-sr-m
spectral radiant intensity of a non-blackbody at wavelength  and temperature , W/m2-sr-m
Notes: Note that the nomenclature or symbology for radiation is not standard across all disciplines or sources. For consistency, we have adopted the naming and symbolic convention used by Howell, J. R. , R. Siegel, and M. P. Mengüç, Thermal Radiation Heat Transfer, Fifth Edition, Taylor & Francis, New York, 2010. This nomenclature also is included in the manual for the radiative function library mentioned in the previous slide.

## Slide 6: Nomenclature - II
total radiant intensity of a blackbody at temperature ,
W/m2-sr-m
total radiant intensity non-blackbody at temperature ,
W/m2-sr-m
integral of spectral radiant intensity of a blackbody at temperature across a finite wavelength band from to , W/m2-sr
integral of spectral radiant intensity of a non-blackbody at temperature across a finite wavelength band from to , W/m2-sr
actual surface temperature, K
ratio temperature, K
measured surface temperature at wavelength  assuming a perfect emitter, K

## Slide 7: Nomenclature - III
total emissivity of a non-blackbody at temperature
wavelength-averaged emissivity for detector
wavelength averaged emissivity ratio for detector 1 and 2,
monochromatic emissivity of a non-blackbody at wavelength 
and temperature
inferred total emissivity of a non-blackbody at temperature
emissivity ratio at two wavelengths and ,
bandwidth of narrow-band detector i, m
equivalent wavelength, , m
wavelength of detector i, m

## Slide 8: Nomenclature - IV
lower wavelength on wide-band radiative thermometer, m
wavelength of maximum spectral intensity, m
upper wavelength on wide-band radiative thermometer, m
ratio of perimeter to diameter for a circle, 3.14159265358979 Stefan-Boltzmann constant, = 5.670401 × 10-8
W/m2-K4
solid angle, sr
Note that the nomenclature or symbology for radiation is not standard across all disciplines or sources. For consistency, we have adopted the naming and symbolic convention used by Howell, J. R. , R. Siegel, and M. P. Mengüç, Thermal Radiation Heat Transfer, Fifth Edition, Taylor & Francis, New York, 2010.

## Slide 9: Introduction
“If there be light, then there is darkness; if cold, then heat; if height, depth also; if solid, then fluid; hardness and softness, roughness and smoothness, calm and tempest, prosperity and adversity, life and death.” ― Pythagoras
Notes: In this introductory section, we will cover the basics of radiation thermometry.

## Slide 10: Radiation Thermometry - I
All bodies above absolute zero emit thermal radiation.
The quantity of power emitted is measured by the radiative intensity.
Energy is distributed across the electromagnetic spectrum and the intensity of emission and the shape of the distribution is dependent on the temperature.
Conversely, the temperature of a body can be measured by the quantity and spectral distribution of energy it emits.
Notes: All bodies above absolute zero emit thermal radiation, which is quantified by the radiative intensity. This energy is distributed across the electromagnetic spectrum and the intensity of emission and the shape of the distribution is dependent on the temperature of the body. Conversely, the temperature of a body can be measured by the quantity and spectral distribution of energy it emits.

## Slide 11: Radiation Thermometry - II
Chart: Wavelength, nm
Radiation thermometry is most often applied to measuring the temperature of heated solid bodies
However, the techniques discussed here are applicable to measuring the temperature of a group of hot particles in a cloud
These applications emit radiation across the entire spectrum (continuum radiation) and we will exclude the measurement of radiation due to electronic or molecular transitions at specific wavelengths (line radiation)
Notes: Measuring the temperature of individual heated solid bodies is the most common application of radiation thermometry, although the techniques discussed here are applicable to measuring the temperature of a group of hot particles in a cloud. These applications emit radiation across the entire spectrum (continuum radiation) and we will exclude the measurement of radiation due to electronic or molecular transitions at specific wavelengths (line radiation).

## Slide 12: What is Temperature?
Temperature is a numerical measure of the internal energy of a body.
Heat flows between bodies at different temperatures.
Heat flows from a body at higher temperature to a body with a lower temperature.
Temperature is measured using a thermometer.
World’s Tallest Thermometer Baker, CA
Notes: Temperature is a numerical measure of the internal energy of a body. Heat flows between bodies at different temperatures and flows from a body at higher temperature to a body with a lower temperature. Temperature is measured using a thermometer.

## Slide 13: Types of Thermometers
Liquid in glass
Thermocouple
Resistance Temperature Devices (RTD)
Thermistors
Radiation
The first four thermometer types are “contact thermometers” in that they require actual physical contact with the object being measured. A radiation thermometer is a “non-contact” measurement device that does not require physical contact with the object to be measured.
Notes: There are many kinds of thermometers. These include simple liquid in glass instruments as well as thermocouples and resistance temperature detectors or RTDs. These are examples of “contact thermometers” in that they require actual physical contact with the object being measured. A radiation thermometer, on the other hand, is a “non-contact” measurement device that does not require physical contact with the object to be measured.

## Slide 14: Advantages of Radiation Thermometers - I
Non-contact. Can be used when the object is:
  Too remote
  Too fragile
  Too hot
  Too sensitive to contamination
Very fast response. Many radiation thermometers do not require the heating of a physical object for measurement, unlike contact devices.
Does not perturb the temperature field. Contact sensors can cause heat losses and distortions to the temperature field resulting in inaccurate measurements.
Immune from electrical interference. Signal transmission is by high-frequency electromagnetic waves not affected by the presence of electrical or magnetic fields.
Notes: There are many advantages to radiation thermometers. One of the most important advantage is that radiation thermometers are non-contact sensors and can be used when the object is too remote, too fragile, too hot, or sensitive to contamination. Radiation thermometers can also have very fast response times since many radiation thermometers do not require the heating of a physical object for measurement. The non-contact nature of radiative thermometry means that the temperature field also is not perturbed by the measurement. Finally radiation thermometers are immune from electrical interference since light is a high-frequency electromagnetic wave and is not affected by the presence of lower frequency electrical or magnetic fields..

## Slide 15: Advantages of Radiation Thermometers - II
For example,
  Temperature measurements can be made for bodies heated above 4000 K where no other physical sensor could survive.
  Response times on the order of nanoseconds are possible with the right detector.
  Temperatures of the sun and distant stars can be determined using the emitted radiation measured on earth.
Notes:   For example, radiation thermometers can make measurements at temperatures above 4000 K where no other physical sensor could survive. One can obtain detectors with response times on the order of nanoseconds with the right detector.
  The temperatures of remote objects such as the sun and distant stars can be determined using the emitted radiation measured on earth.

## Slide 16: Disadvantages of Radiation Thermometers
Do not provide an unambiguous measurement.
  Measurement depends on the optical surface properties, which often are unknown.
More difficult to calibrate
  Measurement is a radiative flux, not a voltage or displacement and calibration standards are more difficult to obtain.
Generally more complex than other measurements methods
  Liquid in glass thermometers or a thermocouple are relatively simple compared to radiation thermometers.
Must have a clear line of site to object being measured.
Notes: There are several disadvantages to radiation thermometers, however. First, radiation thermometers do not provide an unambiguous measurement. The measurement depends on the optical surface properties, which often are unknown. Since the radiation emitted from a body depends not only on the temperature, but also on surface properties, decupling these two effects is often very difficult. Only if the surface is a perfect emitter, can the temperature be determined with certainty. Second, radiation thermometers are more difficult to calibrate. The measurement is a radiative flux, not a voltage or displacement and calibration standards are more difficult to obtain. Radiation thermometers are generally more complex than other measurements methods such as liquid in glass thermometers or thermocouples. Finally, the instrument must have a clear line of site to the object. This may limit their application if the object to be measured is completely enclosed and not visible.

## Slide 17: Radiation Measurement Spectrum
Radiation thermometers measure radiation in the ultra-violet (UV), visible, and infra-red (IR) ranges of the electromagnetic spectrum. These comprise the regions in the spectrum from about 0.1 to 5000 m.
Notes: Radiation thermometers measure radiation in the ultra-violet (UV), visible, and infra-red (IR) ranges of the electromagnetic spectrum. These comprise the regions in the spectrum from about 0.1 to 5000 m.

## Slide 18: Radiation Temperature Measurement
By measuring the radiation intensity or distribution, the temperature of an ideal surface can uniquely be obtained.
This is the basic principle of radiation temperature measurement.
Notes: By measuring the radiation intensity or distribution, the temperature of an ideal surface can uniquely be obtained. This is the basic principle of radiation temperature measurement.

## Slide 19: History of Radiation Thermometry
“Heat cannot be separated from fire, or beauty from The Eternal.” ― Dante Alighieri
Notes: Next, we will cover the history of radiation thermometry.

## Slide 20
On of the first radiation thermometers was the human eye. The human eye can detect a luminance range of 1014, or one hundred trillion (100,000,000,000,000) over an almost 180-degree forward-facing horizontal field of view. The human eye can distinguish about 10 million individual colors.1
1) Judd, Deane B.; Wyszecki, Günter (1975). Color in Business, Science and Industry. Wiley Series in Pure and Applied Optics (third ed.). New York:Wiley-Interscience. p. 388. ISBN 0-471-45212-2.
Notes: The first radiation thermometer was the human eye which has amazing capabilities. The human eye can detect a luminance range of 1014, or one hundred trillion (100,000,000,000,000) while covering an almost 180-degree forward-facing horizontal field of view. The human eye can distinguish about 10 million individual colors.

## Slide 21
As early as 1500 B.C. the process of making iron and steel consistently required that the furnace and the metals be heated to a repeatable temperature. Since there was no method capable of directly measuring temperature, it was discovered that the color of the furnace and the metal could indicate temperature.
Early Roman Blacksmith
Notes: As early as 1500 B.C. the process of making iron and steel consistently required that the furnace and the metals be heated to a repeatable temperature. Since there was no capability to directly measure temperature, it was discovered that the color of the furnace and the metal could roughly indicate temperature.

## Slide 22
Ezekiel (593 to 565 BC) “Then set the empty pot on the coals till it becomes hot and its copper glows, so that its impurities may be melted and its deposit burned away.”
Ezekiel 24:11, New International Version (NIV)
Notes: In the Old Testament, Ezekiel refers to glowing copper as a way to quantify heat and high temperature.

## Slide 23
“…and from the Anvils, the sturdy Strokes in echoing Groans resound the red hot Bars of Steel hiss in the caverns, and the Fire in the Furnace pants.”
P. Virgil Maronis, Aneidos Liber Octavus, 19 B.C., translated by Joseph Davidson, 1890
The Forge of Vulcan by Diego Velázquez, (1630).
Notes: The ancient Roman poet Virgil similarly used the color of red hot steel to describe intense heat and temperature. These examples demonstrate that even the ancients knew that the color and intensity of emitted light from a heated material was an indication of the temperature or quantity of heat in an object.

## Slide 24
Even today, the color of heated steel is used to determine its temperature.
Notes: Even today, the blacksmith uses color of heated steel to determine its temperature as shown in the chart.

## Slide 25: Practical Radiation Thermometers
“Is not Fire a Body heated so hot as to emit Light copiously? For what else is a red hot Iron than Fire? And what else is a burning Coal than red hot Wood?“
— Sir Isaac Newton
Opticks (1704), Book 3, Query 9, 134.
Notes: Next we will discuss the components of practical radiation thermometers including detectors and wavelength selection devices.

## Slide 26: Early Radiation Thermometers
Sketch based on the description from: M’Sweeny, J., “On Regulating the heat of Furnaces by means of Pyrometers, acted on by radiant heat.”, Journal of the Franklin Institute, Vol. III, pp. 13-16, 1829.
Early 1828 M’Sweeny radiation thermometer (pyrometer) used for measuring the temperature of furnaces.
Notes: Possibly the first practical radiation thermometer was developed in 1828 by M’Sweeny who used a mercury in glass thermometer at the end of a copper cone to reflect heat from a furnace. The radiant energy heated the thermometer indicating the temperature of the fire. The principle of radiatively heating a sensor to measure changes in its temperature remains a measurement method today, although semiconductors and electronics have replaced the glass thermometer.

## Slide 27
Disappearing filament pyrometer patent from 1902 which became the industrial and scientific standard pyrometer in the early 20th century
Notes: The disappearing filament pyrometer became the industrial and scientific standard pyrometer in the early 20th century. It was the first device which could measure temperatures above 1000°C.

## Slide 28: How a Disappearing Filament Pyrometer Works
Notes: The disappearing-filament pyrometer measures the temperature of a glowing incandescent object by comparing it to the light of a heated filament. In the design above, a telescope with a thin wire (filament) is placed at the focal plane of the objective lens and is heated by electric current. When seen through the eyepiece, the wire appears silhouetted in front of the hot luminous object under investigation, as shown in the sketch at the bottom left. If the filament is brighter than the background object, then the temperature of the object is below that of the filament. Conversely, if the filament is darker, then the object temperature is higher than the filament, The user adjusts the current through the filament until it seems to "disappear" in front of the glowing object as seen in the sketch on the bottom right. At that point the filament and object are at the same temperature. The user then reads the temperature off the filament current control dial, which is calibrated by the filament's current-vs-temperature curve.

## Slide 29: Modern Development of Detectors
In the times leading up to WWII, electronic detectors were invented that allowed fully automated temperature measurement.
These include thermopile and pyroelectric detectors which actually measure the increase in temperature of a sensor and quantum sensors that measure the current generated when photons strike a semiconductor.
In the last 30 years, solid-state radiation detectors have been developed and refined to provide low-cost solutions for temperature measurement across a wide range of wavelengths and intensities.
Original 1P25 image converter tube developed by RCA in the 1930’s
Notes: While the disappearing filament pyrometer was reasonably accurate and became a standard method for measuring high temperature objects optically, it suffered from one major flaw; measurements required human intervention and depended on the vision of the user to determine the relative brightness of the filament. In the times leading up to WWII, electronic detectors were invented that allowed fully automated temperature measurement. These include thermopile and pyroelectric detectors which actually measure the increase in temperature of a sensor. Quantum sensors, developed later, measure the current generated when photons strike a semiconductor. In the last 30 years, solid-state radiation detectors have been developed and refined to provide low-cost solutions for temperature measurement across a wide range of wavelengths and intensities.

## Slide 30: Three Major Detector Types
Internal photonic
  Electrons are created within a semiconductor material producing a voltage or current
External photonic
  Electrons are emitted from photosensitive material creating a current
Thermal
  Heat is generated that changes the property of a material
Notes: There are three major detector types used in the modern radiation thermometers: internal photonic, external photonic, and thermal detectors.

## Slide 31: Common Radiation Detectors
Photodiode
Uncooled IR detectors
Cryogenically cooled IR detectors
Linear Array Detectors
Thermopile
Photomultiplier Tube
Notes: Pictures of common detectors are shown in the photos. To improve range, many IR detectors are cooled either cryogenically or thermoelectrically to reduce background noise. However, improvements in detectors are rapid, especially in the area of linear and two-dimensional focal plane arrays that allow measurements to be made across the surfaces of objects, not just at one spatial position.

## Slide 32: Wavelength Ranges of Common Detectors
Chart: Wavelength, m
| 0.11 | 1
 | 1 | 1
 | 2 | 2
 | 3 | 3
 | 4 | 4
 | 5 | 5
 | 6 | 6
 | 7 | 7
 | 8 | 8
 | 9 | 9
 | 10 | 10
 | 11 | 11
 | 12 | 12
Notes: The wavelength range of common detectors are shown here. Infrared detectors can be fabricated from semiconductor materials such as Mercury Cadmium Telluride (HgCdTe), Indium Arsenide (InAs), Indium Antimonide (InSb), Lead Sulfide (PbS), and Lead Selenide (PbSe). Visible and near IR detectors are made from Silicon and Germanium, while photomultplier tubes are common for ultraviolet applications. Thermal detectors cover a wide range of wavelengths from the visible to the far infrared. Sensors in the visible to the near IR are generally reserved for measuring higher temperatures, in the range of 1000 K and above. Under the right conditions, longer wavelength detectors such as Mercury Cadmium Telluride can accurately measure surface temperatures of objects at room temperature or below.

## Slide 33: Wavelength Response of Detectors
Notes: The general behavior of thermal and photonic detectors is shown here. Photonic detectors exhibit a response that increases with wavelength until a peak is reached at the cutoff wavelength and then rapidly declines. Thermal detectors exhibit a response that is uniform across the wavelength spectrum.

## Slide 34: Features of Different Detectors
Type | Detector | Features
Internal Photonic | Si Photodiode Ge Photodiode InGaAs PbS PbSe InSb HgCdTe | High speed response Wide spectral range Small size Easy integration
External Photonic | Photomultiplier tube Photo tube Image Tube | High sensitivity High-speed response Large photosensitive area
Thermal | Thermopile Bolometer Pyroelectric | Insensitive to wavelength Slow response
Notes: The characteristics of various detectors are summarized in this slide. For most engineering applications of temperature measurement, internal photonic detectors are usually the best choice. External photonic detectors find use in low-light applications and thermal detectors are sometimes used when low cost is a requirement.

## Slide 35: Ways Detectors are Used
To simplify the analysis and improve measurement accuracy, the light incident on a detector is usually filtered to provide a narrow wavelength or color
The bandwidth is usually made short – 10 to 100 nm wide so that the spectral intensity can be assumed constant across the band
However, this reduces the light input and measured signal
So in cases where high sensitivity is needed, wide-band detectors can be used
Notes: To simplify analysis, the light incident on a detector is usually filtered to provide a narrow wavelength. The spectral intensity can thus be assumed constant across the waveband. Wide-band detectors, which provide a higher signal, can be used, but as we will see later in the lesson, the interpretation of measurements is complicated by the wide bandwidth.

## Slide 36: Methods to Limit Optical Bandwidth - I
Interference Filters
Interferometer
Dispersive Spectrometer
Notes: There are several methods used to limit optical bandwidth. These include interference filters, interferometers, or dispersive spectrometers. The advantages of each are application specific.

## Slide 37: Common interference filter spectral response.
600-nm CW – 12-nw wide Interference Filter
Methods to Limit Optical Bandwidth - II
Notes: The use of interference filters is one of the most common ways to limit the spectral bandwidth to a detector, especially for temperature measurement. The spectral transmission of a typical interference filter is shown here. This filter is designed for a center wavelength at about 596 nm (0.596 µm) and a 12-nm bandwidth. Interference filters across the UV, visible, and IR can be obtained for a nominal cost and the filter bandwidth can be tailored to meet specific requirements. Note that the transmission of the filter, as shown on the plot, is only about 50% even for the most transmissive wavelength. This is true for most components to limit bandwidth. There is always some loss in these components.

## Slide 38
Today, radiation thermometers covering a wide range of wavelengths and temperatures are available from instrument manufacturers.
Additionally, components are readily available to make high precision instruments that can be tailored to meet specific requirements.
Modern Radiation Thermometers
Notes: Today, radiation thermometers covering a wide range of wavelengths and temperatures are available from instrument manufacturers and components are readily available to make high precision instruments that can be tailored to meet specific requirements. In fact, a competent researcher can build a much better radiative thermometer than what is available commercially since most commercial products are targeted for industrial and not laboratory use.

## Slide 39
An engineer uses a readily available handheld radiation thermometer to measure the temperature of a furnace
A precision commercial radiation thermometer for process measurement
Radiation Thermometers in Use
Notes: Here, on the left, we see an engineer using a common handheld radiation thermometer to measure the temperature of a furnace. A precision, commercial radiation thermometer is shown on the right.

## Slide 40: Visible-Wavelength Radiation Thermometer
Detector and Amplifier
Back Optic
Front Optic
Filter
Notes: Shown here is a simplified visible radiation thermometer fabricated from commonly available components. In the visible part of the spectrum, it is possible to use refractive optics as shown here. Light enters from the right and is collimated by a front optic. The light then passes through a filter limiting the bandwidth. The back optic then focuses the light on the active part of a detector. The detector is connected to an amplifier and then measured by voltmeter or other data acquisition device.

## Slide 41: 3-Channel IR Radiation Thermometer
Primary Mirror
Beam Splitters
LN2-Cooled Detectors
Wavelength Filters
To Target
Notes: In this slide we see the conceptual design for a 3-channel, infrared radiation thermometer using three, liquid nitrogen-cooled detectors. The target beam enters from the right and is reflected off a primary mirror and then split into three individual beams. Because the light is in the IR, it more difficult to obtain refractive optical components, so all the optics here are reflective. At the entrance of each detector is an interference filter that provides three different wavelengths or colors for analysis.

## Slide 42: Practical Measurement Tips
To ensure accurate non-contact temperature measurements, be aware of the following:
Field of View
  Ensure the sample area overfills the spot size
Optical Path
  Smoke, fog, dust, etc. can attenuate the signal
Instrument Temperature
  Environments that are too cold or too hot can degrade instrument accuracy
Emissivity
  The surface emissivity can significantly reduce the measured signal
Notes:   To ensure accurate non-contact temperature measurements, the items listed here should be observed whenever possible. Ensure that the sample area overfills the detector to eliminate chances that the detector does not completely image the measured area. Smoke, fog, and dust can attenuate the signal, so it is important that there is clear field of view from the detector to the target. Environments that are too cold or too hot can degrade instrument accuracy. Finally, as we have mentioned, the surface properties of the material also influence the measured temperature. Later in this lesson we will discuss the issue of surface emissivity further.

## Slide 43: Fundamental Physics
“The spectral density of black body radiation ... represents something absolute, and since the search for the absolutes has always appeared to me to be the highest form of research, I applied myself vigorously to its solution.”
— Max Planck
Notes: In this section, we will discuss the fundamental physics underlying radiation thermometry measurements.

## Slide 44: Blackbody - I
A blackbody is an idealized physical body that absorbs all incident electromagnetic radiation and emits electromagnetic radiation called blackbody radiation.
Notes: The first concept we will discuss is that of a blackbody. A blackbody is an idealized physical body that absorbs all incident electromagnetic radiation and emits electromagnetic radiation called blackbody radiation. A blackbody is best visualized by placing it inside a perfectly insulated enclosure of arbitrary shape that is maintained under a vacuum. To maintain equilibrium, the blackbody must emit the same amount of radiation as it absorbs. In order to be consistent with the second law of thermodynamics, the total emitted radiation from a blackbody must increase as the temperature increases.

## Slide 45: Blackbody - II
The radiation emitted from a blackbody is determined by the temperature alone not by the body's shape or composition.
A blackbody in thermal equilibrium has two notable properties:
  It is an ideal emitter: at every frequency and wavelength, it emits as much energy as – or more energy than – any other body at the same temperature.
  It is a diffuse emitter: the energy is radiated isotropically independent of direction.
Notes: The radiation emitted from a blackbody is determined by the temperature alone not by the body's shape or composition. A blackbody in thermal equilibrium has two notable properties:
1) it is an ideal emitter: at every frequency and wavelength, it emits as much energy as – or more energy than – any other body at the same temperature, and 2) it is a diffuse emitter: the energy is radiated isotropically independent of direction
The total radiation emitted from a blackbody is also independent of placement or orientation of the body in the chamber. Furthermore, the blackbody must absorb the same amount of radiation that it emits at each wavelength. The important item to remember, is that a blackbody is a perfect emitter at all wavelengths and temperatures.

## Slide 46: Fundamental Physics - I
All heated bodies emit thermal radiation
The intensity and distribution of energy is governed by Planck’s law and is determined by the wavelength of light and the temperature. For a perfect emitter (blackbody) in a vacuum:
Where is the blackbody spectral radiant intensity at wavelength  and temperature (typically in units of W/m2-sr-m), and and are Planck’s first and second constants, respectively.
 |  | (1)
 |  | (1)
Notes: All heated bodies emit thermal radiation. For a perfect emitter, Max Planck in 1901 quantified the energy emitted by a blackbody as a function of the wavelength of emitted light and the temperature of the body. This relationship is known as Planck’s law.

## Slide 47: Fundamental Physics - II
The intensity distribution is expressed in units of W/m2-sr-m which is a power per area, per unit wavelength, per solid angle.
A solid angle Ω equals 4π times the area subtended by a spherical cap at a given distance divided by the entire area of the sphere at that distance.
The entire sphere has a solid angle of 4π sr.
A
R
Notes: The emissive intensity distribution is expressed in units of Watts per square meter per steridian per micrometer (W/m2-sr-m) which is a power per area, per unit wavelength, per solid angle. A solid angle Ω equals 4π times the area subtended by a spherical cap at a given distance divided by the entire area of the sphere at that distance. The entire sphere has a solid angle of 4π steridians.

## Slide 48
Chart: Emissive Power versus Temperature
| 0.1 | 0.11000000000000001 | 0.12100000000000002 | 0.13310000000000002 | 0.14641000000000004 | 0.16105100000000006 | 0.17715610000000007 | 0.19487171000000009 | 0.21435888100000011 | 0.23579476910000013 | 0.25937424601000014 | 0.28531167061100016 | 0.31384283767210019 | 0.34522712143931022 | 0.37974983358324127 | 0.41772481694156544 | 0.45949729863572203 | 0.50544702849929424 | 0.55599173134922375 | 0.61159090448414621 | 0.67274999493256094 | 0.74002499442581704 | 0.81402749386839879 | 0.89543024325523879 | 0.98497326758076276 | 1.0834705943388392 | 1.1918176537727232 | 1.3109994191499956 | 1.4420993610649953 | 1.5863092971714949 | 1.7449402268886445 | 1.919434249577509 | 2.1113776745352602 | 2.3225154419887866 | 2.5547669861876656 | 2.8102436848064323
3000 | 1.7682e-14 | 8.5912e-13 | 2.8083e-11 | 6.4025e-10 | 1.0519e-08 | 1.2832e-07 | 1.1941e-06 | 8.688e-06 | 5.0538e-05 | 0.00023987 | 0.00094633 | 0.0031557 | 0.0090322 | 0.022499 | 0.049394 | 0.096678 | 0.17047 | 0.2734 | 0.40225 | 0.54725 | 0.69342 | 0.82377 | 0.9231 | 0.98117 | 0.99435 | 0.96536 | 0.90177 | 0.81379 | 0.7121 | 0.60622 | 0.50363 | 0.40943 | 0.32651 | 0.25599 | 0.1977 | 0.15067
3500 | 1.6711e-11 | 4.3555e-10 | 8.0821e-09 | 1.1012e-07 | 1.1332e-06 | 9.0335e-06 | 5.7101e-05 | 0.00029229 | 0.0012351 | 0.0043841 | 0.013281 | 0.034833 | 0.080145 | 0.1637 | 0.30007 | 0.49849 | 0.75728 | 1.0607 | 1.3799 | 1.679 | 1.9227 | 2.084 | 2.1496 | 2.1204 | 2.0092 | 1.8364 | 1.6253 | 1.3978 | 1.1717 | 0.96022 | 0.77119 | 0.60843 | 0.47251 | 0.36186 | 0.27372 | 0.2048
4000 | 2.8486e-09 | 4.6535e-08 | 5.6472e-07 | 5.2302e-06 | 3.7889e-05 | 0.00021954 | 0.0010383 | 0.0040831 | 0.013576 | 0.038752 | 0.096295 | 0.21094 | 0.41204 | 0.72524 | 1.1612 | 1.7058 | 2.3176 | 2.933 | 3.4804 | 3.8961 | 4.1374 | 4.1895 | 4.0643 | 3.794 | 3.4218 | 2.9926 | 2.5465 | 2.1146 | 1.7185 | 1.37 | 1.0738 | 0.82907 | 0.63168 | 0.47565 | 0.35446 | 0.26172 | 0.19167 | 0.13935 | 0.10067
4500 | 1.55e-07 | 1.7607e-06 | 1.5357e-05 | 0.00010534 | 0.00058079 | 0.0026257 | 0.0099103 | 0.031745 | 0.087596 | 0.21106 | 0.44957 | 0.85609 | 1.4723 | 2.3082 | 3.3268 | 4.442 | 5.5337 | 6.4734 | 7.1536 | 7.5085 | 7.5234 | 7.23 | 6.6923 | 5.9903 | 5.2036 | 4.4013 | 3.6355 | 2.9404 | 2.3343 | 1.8229 | 1.4028 | 1.0657 | 0.80043 | 0.59513 | 0.43853 | 0.32058 | 0.23271 | 0.16787 | 0.12042
5000 | 3.7922e-06 | 3.2212e-05 | 0.00021571 | 0.0011637 | 0.0051573 | 0.019117 | 0.060241 | 0.16377 | 0.38928 | 0.81902 | 1.5422 | 2.6255 | 4.0782 | 5.8285 | 7.7233 | 9.5549 | 11.1076 | 12.2051 | 12.7452 | 12.7117 | 12.1653 | 11.2189 | 10.0086 | 8.6686 | 7.3127 | 6.0262 | 4.864 | 3.8543 | 3.0049 | 2.3092 | 1.7521 | 1.3144 | 0.97631 | 0.71878 | 0.52505 | 0.38086 | 0.27457 | 0.19685 | 0.14044
5500 | 5.1878e-05 | 0.0003474 | 0.0018741 | 0.008306 | 0.030789 | 0.097019 | 0.26375 | 0.62693 | 1.319 | 2.4838 | 4.2284 | 6.568 | 9.3869 | 12.4384 | 15.3883 | 17.8889 | 19.6565 | 20.5262 | 20.4723 | 19.5924 | 18.0681 | 16.1189 | 13.9608 | 11.7772 | 9.7052 | 7.8335 | 6.2074 | 4.8394 | 3.719 | 2.8217 | 2.1169 | 1.5724 | 1.1576 | 0.84559 | 0.61339 | 0.4422 | 0.31704 | 0.22619 | 0.16067
 | 0 | 250000000
 | 0 | 250000000
Notes: Here we see a graphical representation of the blackbody emissive intensity as a function of wavelength for various temperatures.

## Slide 49: Planck Function Properties
The Planck function has the following three properties:
The radiant intensity increases with increasing temperature at all wavelengths and for all temperatures (i.e. is always greater than zero).
For any given temperature, the radiant intensity reaches a maximum at a specific wavelength and there are two wavelengths where the radiant intensity is equal for intensities less than the maximum.
The wavelength at the maximum radiant intensity decreases as the temperature increases (Wien’s displacement law).
Notes: As shown in the previous slide, the Planck function has three important properties:
The radiant intensity increases with increasing temperature at all wavelengths and for all temperatures (i.e. is always greater than zero)
For any given temperature, the radiant intensity reaches a maximum at a specific wavelength and there are two wavelengths where the radiant intensity is equal for intensities less than the maximum.
The wavelength at the maximum radiant intensity decreases as the temperature increases. This known as Wien’s law.
As shown in the previous slide, the The Planck function has three important properties:
The radiant intensity increases with increasing temperature at all wavelengths and for all temperatures (i.e. 〖𝑑𝑖_(𝑏,𝜆)〗∕𝑑𝑇 is always zero)
For any given temperature, the radiant intensity reaches a maximum at a specific wavelength and there are two wavelengths where the radiant intensity is equal for intensities less than the maximum.
The wavelength at the maximum radiant intensity decreases as the temperature increases (Wien’s displacement law).

## Slide 50: Wien’s Law
Peak monochromatic emissive power occurs at a wavelength dependent on the temperature
The peak wavelength can be found by differentiating Planck’s law with respect to wavelength
The peak wavelength occurs where:
 |  | (2)
 |  | (2)
Notes: The inverse relationship between the wavelength at the peak emissive intensity and the temperature is known as Wien’s law. Wien’s law can be found by differentiating the Planck equation with wavelength resulting in the equation above.

## Slide 51: Wien’s Law Peak Intensity and Wavelength
Chart: Temperature, K
Notes: The relationship between the peak intensity and the wavelength at the peak intensity versus temperature from Wien’s law is shown here.

## Slide 52: Example 1 - I
E1:What is the emitted spectral intensity for a blackbody at 3000 K and at a wavelength of 0.5 m?
Notes: Now let’s try an example. What is the emitted spectral intensity for a blackbody at 3,000 K and at a wavelength of 0.5 m?

## Slide 53: Example 1 - II
A1: The spectral intensity is calculated from the Planck Equation (Equation 1) using routine bb_ibl from the Excel® library and is:
W/m2-sr-m
Notes: The spectral intensity is calculated from the Planck Equation (Equation 1) using routine bb_ibl from the Excel® library and is equal to W/m2-sr-m.
The spectral intensity is calculated from the Planck Equation (Equation 1) using routine bb_ibl from the Excel library and is equal to 2.60× 〖10〗^5 W/m2-sr-m.

## Slide 54: Example 2 - I
E2: What is emitted power across a 10 nm (0.01 m) waveband centered at 0.5 m and at 3,000 K?
Notes: What is the emitted power across a 10 nm (0.01 m) waveband centered at 0.5 m and at 3,000 K?

## Slide 55: Example 2 - II
A2: Since the bandwidth is small, one can evaluate the total power using the spectral intensity at the given wavelength calculated using routine bb_ibl and then multiplied by the wavelength band:
W/m2-sr-m 0.01 m= W/m2-sr
Notes: Since the bandwidth is small, one can evaluate the total power using the spectral intensity at the given wavelength calculated using routine bb_ibl and then multiplied by the wavelength band. The resulting value is W/m2-sr.
Since the bandwidth is small, one can evaluate the total power using the spectral intensity at the given wavelength calculated using routine bb_ibl and then multiplied by the wavelength band. The value is 2.60× 〖10〗^3 W/m2-sr.

## Slide 56: Response over Wavelength Band - I
The intensity across a wavelength band can be determined by integrating the Planck function:
However, no closed form solution exists, so that the integration must be performed numerically, for example by:
  Polynomial functions (e.g. Simpson’s rule)
  Gaussian quadrature
  Series solution
 |  | (3)
 |  | (3)
Notes: The intensity across a finite wavelength band can be determined by integrating the Planck function. However, no closed form solution for this integral exists, so the integration must be performed numerically, for example by polynomial functions (e.g. Simpson’s rule), Gaussian quadrature, or a series solution.

## Slide 57: Response over Wavelength Band-II
A useful expression for the fraction of power between zero and a specified wavelength is:
So that the fractional power between two wavelengths is:
 |  | (4)
 |  | (4)
 |  | (5)
 |  | (5)
 |  | (6)
 |  | (6)
Notes: A useful expression for the fraction of power between zero and a specified wavelength is given by the series solution above. The fractional power or intensity can be found by subtracting two evaluations of the series solution.

## Slide 58: Fraction of Total Power In-Band Energy
Chart: T, m-K
Notes: The fraction of total in-band energy from zero to a specified wavelength using the series solution is plotted above.

## Slide 59: Response over Wavelength Band-III
Over all wavelengths, the total blackbody intensity is:
which is known as the Stefan-Boltzmann law.
The total hemispherical emissive power is:
The quantity is known as the Stefan-Boltzmann constant and is equal to:
5.6704 ×10-8 W/m2-K4 or 0.1714 ×10-8 Btu/hr-ft2-R4
 |  | (7)
 |  | (7)
 |  | (8)
 |  | (8)
Notes: Integrating over all wavelengths from zero to infinity gives the total blackbody intensity. The relationship that the total blackbody intensity is proportional to the fourth power in temperature is known as the Stefan-Boltzmann law. The proportionality constant between the total hemispherical emissive power and the fourth power of temperature is known as the Stefan-Boltzmann constant and is equal to 5.6704 ×10-8 W/m2-K4.

## Slide 60: Definition of Hemispherical Emissive Power
To calculate the radiant power of a planar surface requires the integration into a hemispherical cap over the surface
is known as the spectral blackbody emissive power
 |  | (9)
 |  | (9)
Notes: The total spectral emission passing through the entire hemisphere in terms of the blackbody intensity requires the integration through all angles of the hemisphere. The blackbody hemispherical emissive power is  times the blackbody intensity.

## Slide 61: Example 3 - I
E3: What is the total emissive power for a body at 1,500 K? What is the emissive power between the wavelengths of 1 to 3 m?
Notes: Example 3. What is the total emitted power for a body at 1,500 K? What is the emitted power between the wavelengths of 1 to 3 m?

## Slide 62: Example 3 - II
A3: The total emissive power is given by Equation 8 and calculated using bb_eb:
2.87 W/m2
The wavelength-temperature products for 1 and 3 m at 1,500K are:
m-K and m-K
0.01285 and = 0.56430
So
W/m2
Chart: T, m-K
Notes: The total emitted power at 1,500 K is given by Equation 8 and calculated using routine bb_eb and is 2.87 W/m2.. The fraction of total energy between 1 and 3 µm is calculated using the series solution for the fraction of power between the bands. The fractional power in band is about 55% and the resulting emitted power is W/m2.
The total emitted power at 1,500 K is given by Equation 8 and calculated using routine bb_eb and is 2.87 × 〖10〗^5 W/m2.. The fraction of total energy between 1 and 3 µm is calculated using the series solution for the fraction of power between the bands. The fractional power in band is about 55% and the resulting emitted power is "1.58"×〖10〗^5 W/m2.

## Slide 63: Fundamental Physics - III
Real surfaces do not emit radiation like an idealized blackbody
Instead, real surfaces emit at a lower rate
The fraction of radiation a surface emits relative to a blackbody is known as the emissivity
The emissivity of a surface varies with direction, wavelength, and temperature
Notes: Unfortunately, real surfaces do not emit radiation like an idealized blackbody, instead, real surfaces emit at a lower rate. The fraction of radiation a surface emits relative to a blackbody is known as the emissivity. The emissivity of a surface varies with the material, its temperature and the direction and wavelength, of the emitted energy.

## Slide 64: Fundamental Physics - IV
Non-black surfaces always emit less energy than a blackbody (Kirchhoff’s Law).
The ratio of the actual energy emitted by a body to that of a blackbody at the same temperature is known as the total hemispherical emissivity and is given by:
Where is the spectral emitted intensity at temperature for an ideal black surface and is the actual emitted intensity of the surface.
 |  | (10)
 |  | (10)
Notes: Non-black surfaces always emit less energy than a blackbody. This is known as Kirchhoff’s Law. The ratio of the actual energy emitted by a body to that of a blackbody at the same temperature is known as the total hemispherical emissivity and is defined by the equation shown.

## Slide 65: Fundamental Physics - V
The emissivity of a surface has a value from 0 to 1.
Surface emissivities vary from near 1 for non-reflective materials such as graphite and carbon black to values less than 0.1 for reflective surfaces such as polished gold and copper.
The special case where the emissivity is constant across the wavelength band is known as a gray surface.
Notes: The emissivity of a surface has a value from 0 to 1. Surface emissivities vary from near 1 for non-reflective materials such as graphite and carbon black to values less than 0.1 for reflective surfaces such as polished gold and copper. The special case where the emissivity is constant across the wavelength band is known as a gray surface.

## Slide 66: Room-Temperature Emissivities of Common Materials
Metals |  | Non-Metals | 
Material | Emissivity | Material | Emissivity
Aluminum Commercial Sheet | 0.09 | Asphalt | 0.93
Aluminum Highly Polished | 0.039 - 0.057 | Black Epoxy Paint | 0.89
Beryllium | 0.18 | Brick, red rough | 0.93
Brass Polished | 0.03 | Carbon, not oxidized | 0.81
Cast Iron, newly turned | 0.44 | Clay | 0.91
Copper electroplated | 0.03 | Concrete | 0.85
Copper Nickel Alloy, polished | 0.059 | Cotton cloth | 0.77
Copper Polished | 0.023 - 0.052 | Glass smooth | 0.92 - 0.94
Gold not polished | 0.47 | Ice rough | 0.985
Gold polished | 0.025 | Ice smooth | 0.966
Inconel X Oxidized | 0.71 | Lampblack paint | 0.96
Iron, dark gray surface | 0.31 | Magnesium Oxide | 0.20 - 0.55
Lead Oxidized | 0.43 | Masonry Plastered | 0.93
Lead pure unoxidized | 0.057 - 0.075 | Paper | 0.93
Magnesium Polished | 0.07 - 0.13 | Plaster | 0.98
Mercury liquid | 0.1 | Plaster board | 0.91
Molybdenum polished | 0.05 - 0.18 | Plastics | 0.91
Nickel, elctroplated | 0.03 | Pyrex | 0.92
Platinum, polished plate | 0.054 - 0.104 | Rubber, natural hard | 0.91
Silver Polished | 0.02 - 0.03 | Salt | 0.34
Stainless Steel, polished | 0.075 | Sand | 0.76
Steel Polished | 0.07 | Silica | 0.79
Tin unoxidized | 0.04 | Silicon Carbide | 0.83 - 0.96
Titanium polished | 0.19 | Soil | 0.38
Tungsten aged filament | 0.032 - 0.35 | Teflon | 0.92
Tungsten polished | 0.04 | Water | 0.95 - 0.963
Wrought Iron | 0.94 | Wood Oak, planned | 0.885
Zinc polished | 0.045 | Wood, Pine | 0.95
Selected values taken from The Engineering Toolbox, Emissivity Coefficients of some common Materials, http://www.engineeringtoolbox.com/emissivity-coefficients-d_447.html
Notes: Shown above is a tabulation of the emissivities for many common materials. Again, metals possess low emissivities while electrical insulators exhibit high emissivities.

## Slide 67: Fundamental Physics - VI
The spectral emissivity is temperature and wavelength dependent and can vary substantially as the surface conditions change.
There is no general theory that allows one to calculate the emissivity for all materials, so values are most often obtained experimentally.
Notes: The spectral emissivity is temperature and wavelength dependent and can vary substantially as the surface conditions change. There is no general theory that allows one to calculate the emissivity for all materials (although some specialized theories work in limited cases, for example for metals). Emissivity values are most often obtained experimentally.

## Slide 68: Spectral Variation of Emissivity for Graphite, Carbon-Carbon, and Carbon-Silicon Carbide
Data from Neuer, G., “Emissivity measurements on graphite and composite materials in the visible and infrared spectral range,” QIRT 92, Eurotherm Series 27, EETI ed., Paris, 1992.
Chart: Wavelength, m
| 0.61006089199864399 | 0.64901030514009705 | 0.72467729676769299 | 0.79838657596113805 | 0.88890689861032501 | 0.96355617418722195 | 1.29834450345702 | 2.1005224010755401 | 3.4998342852116799 | 4.3161428397873802 | 6.3040558741830397
Graphite | 0.92954 | 0.91501 | 0.89616 | 0.87409 | 0.88378 | 0.87786 | 0.87786 | 0.84933 | 0.83856 | 0.8251 | 0.82241
C/SiC | 0.87732 | 0.87517 | 0.8784 | 0.89132 | 0.8924 | 0.85686 | 0.86602 | 0.77342 | 0.71258 | 0.65552
C/C | 0.76265 | 0.74166 | 0.75188 | 0.75188 | 0.76265 | 0.79334 | 0.76319 | 0.74004 | 0.70128 | 0.65121 | 0.60222
 | 0 | 1
 | 0 | 1
 | 0 | 1
 | 0 | 1
 | 0 | 1
 | 0 | 1
 | 0 | 1
 | 0 | 1
 | 0 | 1
Notes: The plot shows the variation of the emissivity for graphite and silicon carbide with temperature. Note the drop-off in emissivity as the wavelength increases.

## Slide 69: Emissivity as a Function of Temperature
Total normal emissivity of graphite, carbon/carbon, and C/SiC composite. From Neuer, G., Emissivity Measurements on graphite and composite materials in the visible, and infrared spectral range, QIRT 92-Eurotherm Series 27 – EETI, Paris, 1992
Chart: Temperature, K
| 1146.2346244094399 | 1155.84140913538 | 1279.3038500794401 | 1352.8225848879899 | 1431.1236676241799 | 1448.5011100939601 | 1475.5326872691601 | 1520.02588465545 | 1554.90177509811 | 1555.62780811683 | 1640.15804370929 | 1641.0997821900901 | 1635.07581257825 | 1673.7712680324501 | 1757.8648315918999 | 1815.05834569693 | 1846.02523227795 | 1894.4958279406901 | 1952.6416027441901 | 2026.2971263823499
Graphite | 0.76838 | 0.77409 | 0.80884 | 0.81298 | 0.8216 | 0.82528 | 0.83101 | 0.8339 | 0.83189 | 0.85064 | 0.82666 | 0.8287 | 0.84541 | 0.84748 | 0.85733 | 0.85126 | 0.8521 | 0.84684 | 0.842 | 0.83554
C/C | 0.72847 | 0.75503 | 0.77303 | 0.77222 | 0.77672 | 0.79146 | 0.80048 | 0.80296
C/SiC | 0.65188 | 0.65764 | 0.66544 | 0.66387 | 0.67003 | 0.69907 | 0.70114 | 0.73789 | 0.73917 | 0.73228 | 0.73025 | 0.73398 | 0.74951
Notes: Here we show the variation of emissivity with temperature. For graphite and silicon carbide, the data shows that the emissivity increases with increasing temperature.

## Slide 70: Methods to Increase Emissivity
It is possible to increase the emissivity if physical modification to the sample surface is allowed. The following methods increase the effective emissivity of the sample to values near 1 and therefore decrease the relative uncertainty in derived temperatures.
High Emissivity Paints and Coatings ( > 0.95)
  Organic and silicon based brush on coatings are good up to 625 K
  Ceramic-based, black-pigmented coating for ceramic fiber modules and refractories are good to 1,650 K
Drilled cavity
  Drill a hole into the sample and measure the temperature inside the hole
“Volcano” cavity
  Build up a “volcano” cavity on the surface using ceramic or carbon cements
Notes:   It is possible to increase the emissivity to a value close to one and therefore decrease the uncertainty in the emissivity if physical modification to the sample surface or configuration is allowed. High emissivity paints and coating can be used to increase the surface emissivity to a value greater than 0.95. Organic and silicon based coatings can be used up to around 625 K while ceramic based coatings are good to up to around 1,650 K. The effective surface emissivity can be increased by physical modifications as well. Fabricating a cavity on the sample and measuring the temperature in the cavity creates a geometry emulating a blackbody. The cavity can be made by drilling a hole into the sample or by building up a “volcano” shape using high-temperature ceramic or graphite cements. Later, we will see why increasing the surface emissivity is beneficial in radiative temperature measurement applications.

## Slide 71: Summary of Part I
All bodies above absolute zero emit thermal radiation.
The quantity of power emitted is measured by the radiative intensity.
Energy is distributed across the electromagnetic spectrum and the intensity of emission and the shape of the distribution is dependent on the temperature.
Conversely, the temperature of a body can be measured by the quantity and spectral distribution of energy it emits.
A blackbody is an idealized physical body that absorbs all incident electromagnetic radiation and emits electromagnetic radiation called blackbody radiation.
The intensity and distribution of energy is governed by Planck’s law and is determined by the wavelength of light and the temperature.
Real surfaces do not emit radiation like an idealized blackbody, but instead, real surfaces emit at a lower rate.
The fraction of radiation a surface emits relative to a blackbody is known as the emissivity.
Notes: To summarize what we have learned in Part I of this lesson: All bodies above absolute zero emit thermal radiation. The quantity of power emitted is measured by the radiative intensity. Energy is distributed across the electromagnetic spectrum and the intensity of emission and the shape of the distribution is dependent on the temperature. Conversely, the temperature of a body can be measured by the quantity and spectral distribution of energy it emits.
A blackbody is an idealized physical body that absorbs all incident electromagnetic radiation and emits electromagnetic radiation called blackbody radiation.
The intensity and distribution of energy emitted by a blaclbody is governed by Planck’s law and is determined by the wavelength of light and the temperature.
Real surfaces do not emit radiation like an idealized blackbody, but instead, real surfaces emit radiation at a lower rate.
The fraction of radiation a surface emits relative to a blackbody is known as the emissivity.

## Slide 72: This is the end of Part I. Please continue to Part II.
Notes: This is the end of Part I. Please continue to Part II.