## Slide 1: Understanding Radiation Thermometry – Part II
Timothy K. Risch
NASA Armstrong Flight Research Center
November 27, 2015
Narrated by Nicholas DeStefano
Notes: This the second of a two-part NASA NESC Academy lesson on radiative thermometry. In this part we will learn about modern measurement methods in radiative thermometry and provide some applications. As in Part I, many documented examples are used to illustrate the concepts that are presented throughout this lesson.

## Slide 2: Lesson Plan
Part I
End off Part 1
Nomenclature
Introduction
History of Radiation Thermometry
Fundamental Physics
Practical Radiation Thermometers
Practical Measurement Techniques
Calibration
Final Thoughts
For Further Reading
Part II
Notes: In this part of the lesson we will discuss practical radiation thermometers and measurement techniques. We will then review calibration techniques and then end with a summary of the lesson and some final thoughts on the subject.

## Slide 3: Preliminaries
An Excel® worksheet containing solutions to all examples are available from Tim Risch at Timothy.K.Risch@NASA.gov
Excel®, Visual Basic for Applications®, and Microsoft® are registered trademarks of Microsoft Corporation.
Notes: All the examples used in this lesson are contained in an Excel® file developed by the author. This file utilizes a library of radiative functions written in Visual Basic for Applications® (VBA) that is also available. To obtain the Excel® file and library, please contact Tim Risch at Timothy.K.Risch@NASA.gov.

## Slide 4: Nomenclature - I
speed of light, 2.99792458 × 108 m/s
Planck’s first constant, = 1.191043 × 108 W-m4/m2-sr
Planck’s second constant, = 14,387.75 m-K
Constant in Wien’s displacement law
Constant in equation for maximum blackbody intensity
detector spectral response function for detector I
fraction of spectral intensity between two wavelength- temperature values and
Planck’s constant, 6.626068 × 10-34 J-s
Boltzmann constant, 1.3806503 × 10-23 J/K
spectral emissive intensity of a perfect blackbody at wavelength  and temperature , W/m2-sr-m
spectral radiant intensity of a non-blackbody at wavelength  and temperature , W/m2-sr-m
Notes: Note that the nomenclature or symbology for radiation is not standard across all disciplines or sources. For consistency, we have adopted the naming and symbolic convention used by Howell, J. R. , R. Siegel, and M. P. Mengüç, Thermal Radiation Heat Transfer, Fifth Edition, Taylor & Francis, New York, 2010. This nomenclature also is included in the manual for the radiative function library mentioned in the previous slide.

## Slide 5: Nomenclature - II
total radiant intensity of a blackbody at temperature ,
W/m2-sr-m
total radiant intensity non-blackbody at temperature ,
W/m2-sr-m
integral of spectral radiant intensity of a blackbody at temperature across a finite wavelength band from to , W/m2-sr
integral of spectral radiant intensity of a non-blackbody at temperature across a finite wavelength band from to , W/m2-sr
actual surface temperature, K
ratio temperature, K
measured surface temperature at wavelength  assuming a perfect emitter, K

## Slide 6: Nomenclature - III
total emissivity of a non-blackbody at temperature
wavelength-averaged emissivity for detector
wavelength averaged emissivity ratio for detector 1 and 2,
monochromatic emissivity of a non-blackbody at wavelength 
and temperature
inferred total emissivity of a non-blackbody at temperature
emissivity ratio at two wavelengths and ,
bandwidth of narrow-band detector i, m
equivalent wavelength, , m
wavelength of detector i, m

## Slide 7: Nomenclature - IV
lower wavelength on wide-band radiative thermometer, m
wavelength of maximum spectral intensity, m
upper wavelength on wide-band radiative thermometer, m
ratio of perimeter to diameter for a circle, 3.14159265358979 Stefan-Boltzmann constant, = 5.670401 × 10-8
W/m2-K4
solid angle, sr
Note that the nomenclature or symbology for radiation is not standard across all disciplines or sources. For consistency, we have adopted the naming and symbolic convention used by Howell, J. R. , R. Siegel, and M. P. Mengüç, Thermal Radiation Heat Transfer, Fifth Edition, Taylor & Francis, New York, 2010.

## Slide 8: Practical Measurement Techniques
“To point out of the advantages that would arise from ascertaining the heat of a body at a very high temperature would be unnecessary, the importance of the subject is allowed.”
- J. M’Sweeny, M.D. - 1829
Notes: In this section we will cover practical measurement techniques for determining the temperature of heated solid bodies. We will discuss measurement techniques that use single detectors and also multiple detectors. We will also quantify the uncertainty in the measurements and provide guidelines to ensure accurate results.

## Slide 9: Radiation Temperature Measurement - I
For a blackbody, there are three primary ways to determine the temperature by measuring:
  The total emitted radiation
  The distribution across wavelengths or the total radiation across a wavelength band
  The emitted radiation at one wavelength
Measuring any of these should uniquely determine the temperature
Notes: For a blackbody, there are three primary ways to determine the temperature by measuring: 1) The total emitted radiation, or 2) measuring the emitted radiation distribution across wavelengths or the total radiation across a wavelength band, or 3) measuring the emitted radiation at one wavelength. Any one of these methods should uniquely determine the temperature.

## Slide 10: Radiation Temperature Measurement - II
As discussed before, real surfaces with emissivities less than one do not behave like a blackbody.
Instead, they emit radiation at rates less than a blackbody
This introduces a new problem, for any given measurement of a non-ideal surface, we now have two unknowns:
  Temperature
  Emissivity
Without assuming either one or the other, the state of the surface cannot be uniquely determined.
Notes: As we discussed before, real surfaces with emissivities less than one do not behave like a blackbody, instead, they emit radiation at rates less than a blackbody. This introduces a new problem, for any given measurement of a non-ideal surface, we now have two unknowns: temperature and emissivity. Without assuming either one or the other, the state of the surface cannot be uniquely determined.

## Slide 11: Three Possible Solutions:
For a single measurement, assume the emissivity and calculate the temperature (spectral method).
Make two measurements at different wavelengths and assume a relationship between the emissivities at each wavelength and calculate a single temperature (ratio method).
Make multiple measurements at different wavelengths and assume some functional form of the emissivity and find the best fit to the temperature and emissivity (multi-spectral method).
For now, all of the examples assume narrow-band detectors, but later we show how to accommodate wide-band detectors.
Notes: There are three possible solutions to determine the surface temperature from radiation measurements: 1) for a single measurement, assume the emissivity and calculate the temperature (spectral method), 2) make two measurements at different wavelengths and assume a relationship between the emissivities at each wavelength and calculate a single temperature (ratio method), or 3) make multiple measurements at different wavelengths and assume some functional form of the emissivity. In this case we then find the best fit to the temperature and emissivity together (multi-spectral method). For now, all of the examples assume narrow-band detectors, but later we show how to accommodate wide-band detectors

## Slide 12: Some Temperature Definitions
“True” Surface Temperature - The actual temperature of the surface; what we would measure with a thermometer not affected by the surface emissivity. This is the temperature we would measure if, for example, we used a contact thermometer such as a thermocouple.
“Equivalent Blackbody” Surface Temperature - The temperature we would infer from a radiation thermometer if we assumed that the surface was a perfect emitter. Again, because all real surfaces emit less than a blackbody, the equivalent blackbody surface temperature will always be lower than the true surface temperature.
Notes: Before we proceed, let us provide some temperature definitions, and clearly define what we mean by “equivalent blackbody” surface temperature and “true” surface temperature. The “true” surface temperature is the actual temperature of the surface; what we would measure with a thermometer not affected by the surface emissivity. In other words, this is the temperature we would measure if, for example, we used a contact sensor such as a thermocouple. The “equivalent blackbody” surface temperature is the temperature we would infer from a radiation thermometer if we assumed that the surface was a blackbody or perfect emitter. Again, because all real surfaces emit less than a blackbody, the equivalent blackbody surface temperature will always be lower than the true surface temperature.

## Slide 13: Method 1 – Spectral Method - I
From a single equivalent blackbody temperature measurement at a specific wavelength , assume the spectral emissivity and calculate the “true” surface temperature . This can be achieved by using the definitions of black and gray emissive intensity:
And then solving for the true surface temperature from:
 |  | (11)
 |  | (11)
 |  | (12)
 |  | (12)
Notes: First, the spectral method. From a single equivalent blackbody temperature measurement at a specific wavelength , assume the spectral emissivity and calculate the “true” surface temperature . This can be achieved by using the definitions of black and gray emissive intensity as shown and then solving for the true surface temperature.
First, the spectral method. From a single equivalent blackbody temperature measurement 𝑇_𝜆, assume the spectral emissivity ε_𝜆 and calculate the “true” surface temperature 𝑇. This can be achieved by using the definitions of black and gray emissive intensity as shown and then solving for the true surface temperature.

## Slide 14: Method 1 – Spectral Method - II
Solving for true surface temperature this is:
Equation 13 can be simplified under the special case when (Wien’s approximation) then:
 |  | (13)
 |  | (13)
 |  | (14)
 |  | (14)
Notes: The true surface temperature is given by Equation 13 in terms of the spectral emissivity at wavelength and the measured equivalent blackbody temperature . In the special case where the product of the wavelength and temperature are much less than Planck’s second constant , or 14,388 µm-K then Equation 13 can be simplified to the expression shown in Equation 14. The special case where the wavelength-temperature product is much less than is known as Wien’s approximation.
The true surface temperature 𝑇 is given by equation 13 in terms of the spectral emissivity ε_𝜆 at wavelength 𝜆 and the measured equivalent blackbody temperature 𝑇_𝜆. In the special case where the product of the wavelength and temperature are much less than Planck’s second constant 𝐶_2, or 14,388 µm-K then equation 13 can be simplified to the expression shown in Equation 14. The special case where the wavelength-temperature product is much less than 𝐶_2 is known as Wien’s approximation.

## Slide 15: Method 1 – Spectral Method - IV
Chart: Emissivity
| 0.1 | 0.12000000000000001 | 0.14000000000000001 | 0.16 | 0.18 | 0.19999999999999998 | 0.21999999999999997 | 0.23999999999999996 | 0.25999999999999995 | 0.27999999999999997 | 0.3 | 0.32 | 0.34 | 0.36000000000000004 | 0.38000000000000006 | 0.40000000000000008 | 0.4200000000000001 | 0.44000000000000011 | 0.46000000000000013 | 0.48000000000000015 | 0.50000000000000011 | 0.52000000000000013 | 0.54000000000000015 | 0.56000000000000016 | 0.58000000000000018 | 0.6000000000000002 | 0.62000000000000022 | 0.64000000000000024 | 0.66000000000000025 | 0.68000000000000027 | 0.70000000000000029 | 0.72000000000000031 | 0.74000000000000032 | 0.76000000000000034 | 0.78000000000000036 | 0.80000000000000038 | 0.8200000000000004 | 0.84000000000000041 | 0.86000000000000043 | 0.88000000000000045 | 0.90000000000000047 | 0.92000000000000048 | 0.9400000000000005 | 0.96000000000000052 | 0.98000000000000054 | 1.0000000000000004
500 | 1.087 | 1.0795 | 1.0733 | 1.068 | 1.0634 | 1.0592 | 1.0555 | 1.0522 | 1.0491 | 1.0463 | 1.0437 | 1.0412 | 1.039 | 1.0368 | 1.0348 | 1.0329 | 1.0311 | 1.0294 | 1.0277 | 1.0262 | 1.0247 | 1.0233 | 1.0219 | 1.0206 | 1.0193 | 1.0181 | 1.0169 | 1.0158 | 1.0147 | 1.0136 | 1.0126 | 1.0115 | 1.0106 | 1.0096 | 1.0087 | 1.0078 | 1.0069 | 1.0061 | 1.0053 | 1.0045 | 1.0037 | 1.0029 | 1.0022 | 1.0014 | 1.0007 | 1
1000 | 1.1905 | 1.1728 | 1.1583 | 1.146 | 1.1353 | 1.126 | 1.1176 | 1.1101 | 1.1033 | 1.0971 | 1.0913 | 1.086 | 1.0811 | 1.0764 | 1.0721 | 1.068 | 1.0642 | 1.0605 | 1.0571 | 1.0538 | 1.0506 | 1.0476 | 1.0447 | 1.042 | 1.0394 | 1.0368 | 1.0344 | 1.032 | 1.0297 | 1.0275 | 1.0254 | 1.0234 | 1.0214 | 1.0194 | 1.0176 | 1.0158 | 1.014 | 1.0123 | 1.0106 | 1.009 | 1.0074 | 1.0058 | 1.0043 | 1.0028 | 1.0014 | 1
2000 | 1.4687 | 1.4164 | 1.3749 | 1.3408 | 1.3122 | 1.2875 | 1.266 | 1.247 | 1.2299 | 1.2146 | 1.2006 | 1.1879 | 1.1761 | 1.1653 | 1.1552 | 1.1458 | 1.1369 | 1.1287 | 1.1209 | 1.1135 | 1.1065 | 1.0999 | 1.0936 | 1.0876 | 1.0818 | 1.0764 | 1.0711 | 1.0661 | 1.0612 | 1.0566 | 1.0521 | 1.0478 | 1.0436 | 1.0396 | 1.0357 | 1.032 | 1.0283 | 1.0248 | 1.0214 | 1.0181 | 1.0149 | 1.0117 | 1.0087 | 1.0057 | 1.0028 | 1
4000 | 2.3745 | 2.1671 | 2.0135 | 1.8943 | 1.7987 | 1.7198 | 1.6533 | 1.5965 | 1.5471 | 1.5037 | 1.4652 | 1.4307 | 1.3996 | 1.3714 | 1.3456 | 1.3219 | 1.3001 | 1.2799 | 1.2612 | 1.2437 | 1.2273 | 1.2119 | 1.1975 | 1.1838 | 1.1709 | 1.1587 | 1.1471 | 1.1361 | 1.1256 | 1.1156 | 1.1061 | 1.0969 | 1.0882 | 1.0798 | 1.0717 | 1.064 | 1.0565 | 1.0493 | 1.0424 | 1.0357 | 1.0293 | 1.023 | 1.017 | 1.0112 | 1.0055 | 1
8000 | 4.4063 | 3.802 | 3.3682 | 3.0412 | 2.7854 | 2.5796 | 2.4103 | 2.2683 | 2.1475 | 2.0433 | 1.9525 | 1.8725 | 1.8015 | 1.738 | 1.6808 | 1.6291 | 1.582 | 1.5389 | 1.4993 | 1.4628 | 1.429 | 1.3977 | 1.3685 | 1.3412 | 1.3157 | 1.2917 | 1.2691 | 1.2479 | 1.2278 | 1.2088 | 1.1908 | 1.1737 | 1.1574 | 1.1419 | 1.1271 | 1.113 | 1.0995 | 1.0866 | 1.0742 | 1.0623 | 1.0509 | 1.0399 | 1.0294 | 1.0192 | 1.0094 | 1
500 | 1.087 | 1.0795 | 1.0733 | 1.068 | 1.0634 | 1.0592 | 1.0555 | 1.0522 | 1.0491 | 1.0463 | 1.0437 | 1.0412 | 1.039 | 1.0368 | 1.0348 | 1.0329 | 1.0311 | 1.0294 | 1.0277 | 1.0262 | 1.0247 | 1.0233 | 1.0219 | 1.0206 | 1.0193 | 1.0181 | 1.0169 | 1.0158 | 1.0147 | 1.0136 | 1.0126 | 1.0115 | 1.0106 | 1.0096 | 1.0087 | 1.0078 | 1.0069 | 1.0061 | 1.0053 | 1.0045 | 1.0037 | 1.0029 | 1.0022 | 1.0014 | 1.0007 | 1
1000 | 1.1905 | 1.1728 | 1.1583 | 1.146 | 1.1353 | 1.126 | 1.1176 | 1.1101 | 1.1033 | 1.0971 | 1.0913 | 1.086 | 1.0811 | 1.0764 | 1.0721 | 1.068 | 1.0642 | 1.0605 | 1.0571 | 1.0538 | 1.0506 | 1.0476 | 1.0447 | 1.042 | 1.0394 | 1.0368 | 1.0344 | 1.032 | 1.0297 | 1.0275 | 1.0254 | 1.0234 | 1.0214 | 1.0194 | 1.0176 | 1.0158 | 1.014 | 1.0123 | 1.0106 | 1.009 | 1.0074 | 1.0058 | 1.0043 | 1.0028 | 1.0014 | 1
2000 | 1.4708 | 1.4179 | 1.3761 | 1.3418 | 1.313 | 1.2882 | 1.2666 | 1.2475 | 1.2304 | 1.215 | 1.201 | 1.1882 | 1.1764 | 1.1655 | 1.1554 | 1.146 | 1.1371 | 1.1288 | 1.121 | 1.1136 | 1.1066 | 1.1 | 1.0937 | 1.0877 | 1.0819 | 1.0764 | 1.0712 | 1.0661 | 1.0613 | 1.0566 | 1.0522 | 1.0478 | 1.0437 | 1.0397 | 1.0358 | 1.032 | 1.0284 | 1.0248 | 1.0214 | 1.0181 | 1.0149 | 1.0117 | 1.0087 | 1.0057 | 1.0028 | 1
4000 | 2.7789 | 2.4358 | 2.2056 | 2.0387 | 1.9111 | 1.8098 | 1.727 | 1.6577 | 1.5987 | 1.5478 | 1.5031 | 1.4637 | 1.4284 | 1.3967 | 1.368 | 1.3418 | 1.3178 | 1.2957 | 1.2753 | 1.2564 | 1.2387 | 1.2222 | 1.2067 | 1.1922 | 1.1785 | 1.1655 | 1.1533 | 1.1416 | 1.1306 | 1.1201 | 1.1101 | 1.1005 | 1.0914 | 1.0826 | 1.0742 | 1.0661 | 1.0584 | 1.0509 | 1.0438 | 1.0368 | 1.0302 | 1.0237 | 1.0175 | 1.0115 | 1.0056 | 1
8000 | 21.4947 | 9.5142 | 6.325 | 4.843 | 3.9842 | 3.4224 | 3.0252 | 2.7289 | 2.4991 | 2.3152 | 2.1645 | 2.0387 | 1.9318 | 1.8399 | 1.7599 | 1.6895 | 1.6271 | 1.5713 | 1.5212 | 1.4758 | 1.4345 | 1.3967 | 1.362 | 1.33 | 1.3005 | 1.273 | 1.2474 | 1.2235 | 1.2011 | 1.1801 | 1.1603 | 1.1416 | 1.124 | 1.1074 | 1.0915 | 1.0765 | 1.0622 | 1.0486 | 1.0356 | 1.0232 | 1.0114 | 1
Notes: In this figure, we plot the temperature ratio as a function of emissivity for several wavelength product values. Curves for both the full Planck expression (Equation 13) and Wien’s approximation (Equation 14) are shown. Note while the temperature ratio using Wien’s law approaches that of the Planck equation at small values of , at large values and low emissivities the results from the two equations diverge.
In this figure, we plot the temperature ratio 𝑇/𝑇_𝜆using the full Planck expression (Equation 13) and Wien’s approximation (Equation 14). Note that at large values of 〖𝜆𝑇〗_𝜆, the two equations diverge.

## Slide 16: Example 4 - I
E4: What is the true surface temperature when a surface at an equivalent blackbody temperature of 3,820 K is measured with a 0.5-m wavelength detector and an emissivity of 0.8 is assumed?
Notes: Example 4. What is the true surface temperature when a surface at an equivalent blackbody temperature of 3,820 K is measured with a 0.5-m wavelength detector and an emissivity of 0.8 is assumed?
Example 4. What is the true surface temperature when a surface at an equivalent blackbody temperature of 3,820 K is measured with at 0.5-𝜇m wavelength detector and an emissivity of 0.8 is assumed.

## Slide 17: Example 4 - II
A4: The wavelength temperature product is 3,820 K times 0.5 m or 1,910 m-K. This is much less than or 14,388 m-K. So, Wien’s approximation is valid and routine bb_tstw can be used to give:
Note that this value is within about of the value calculated with the full Planck equation.
Notes: The wavelength temperature product is 3,820 K times 0.5 m or 1,910 m-K. This is much less than or 14,388 m-K. So, Wien’s approximation is valid and Equation 14 can be used. The true surface temperature is calculated to be K. Note that this value is within about of the value calculated with the full Planck equation.
The wavelength temperature product is 3,820 K times 0.5 𝜇m or 1,910 𝜇m-K. This is much less than 𝐶_2 or 14,388 m-K. So, Wien’s approximation is valid and Equation 14 can be used. The true surface temperature is calculated to be 3,937 K.

## Slide 18: Example 5 - I
E5: Repeat Example 4, except with an 8-m detector. What is the true surface temperature?
Notes: Example 5. Repeat Example 4, except with an 8-m detector. What is the true surface temperature?
Example 5. Repeat Example 4, except with an 8-𝜇m detector. What is the true surface temperature?

## Slide 19: Example 5 - II
A5: The wavelength temperature product is 3,820 K times 8 m or 30,560 m-K. This is greater than or 14,388 m-K. So, the full Planck equation (Equation 1) must be used using routine bb_tst:
If indeed we had used Wien’s approximation, the calculated true surface temperature would have been over 7,200 K and off by over 2,600 K.
Notes: The wavelength temperature product is 3,820 K times 8 m or 30,560 m-K. This is greater than or 14,388 m-K. So, the full Planck equation (Equation 1) must be used using routine bb_tst. The calculated temperature is . If indeed we had used Wien’s approximation, the calculated true surface temperature would have been over 7,200 K and off by over 2,600 K.
The wavelength temperature product is 3,820 K times 8 𝜇m or 30,560 𝜇m-K. This is greater than 𝐶_2 or 14,388 m-K. So, the full Planck equation (Equation 1) must be used using routine bb_tst. The calculated temperature is 4,579 K. If indeed we had used Wien’s approximation, the calculated true surface temperature would have been over 7,200 K and off by over 2,600 K.

## Slide 20: Method 1 – Spectral Method - V
We can see by looking at Equation 13:
that the correction to the measured surface temperature is dependent on the wavelength and when Wien’s approximation applies in Equation 14:
the correction is proportional to the wavelength. This suggests that using a short wavelength detector will minimize the effect of emissivity on our temperature measurement.
 |  | (13)
 |  | (13)
 |  | (14)
 |  | (14)
Notes: We can see by looking at Equation 13, that the correction to the measured surface temperature is dependent on the wavelength and when Wien’s approximation applies in Equation 14, the correction is proportional to the wavelength. This suggests that using a short wavelength detector will minimize the effect of emissivity on our temperature measurement.

## Slide 21: Spectral Method Error Estimation - I
By differentiating Equation 13 with respect to emissivity, we can determine the sensitivity of the computed “true” surface temperature to the assumed emissivity for a detector with wavelength . The result is:
applying Wien’s approximation, this becomes:
 |  | (15)
 |  | (15)
 |  | (16)
 |  | (16)
Notes: By differentiating Equation 11 with respect to emissivity, we can determine the sensitivity of the computed “true” surface temperature to the assumed emissivity for a detector with wavelength . The result is shown in Equation 15. If Wien’s approximation applies, then the sensitivity of the surface temperature to emissivity becomes proportional to the wavelength-temperature product, as shown in Equation 16.
By differentiating equation 11 with respect to emissivity, we can determine the sensitivity of the computed the “true” surface temperature to the assumed emissivity for a detector 𝑖 with wavelength 𝜆_𝑖. The result is shown in Equation 15. If Wien’s approximation applies, then sensitivity of the surface temperature to emissivity becomes proportional to the wavelength-temperature product.

## Slide 22: Spectral Method Error Estimation - II
Using the sensitivity and an estimate of the uncertainty in the emissivity produces an estimate of the uncertainty in the surface temperature:
 |  | (17)
 |  | (17)
Notes: Using the sensitivity and an estimate of the uncertainty in the emissivity produces an estimate of the uncertainty in the surface temperature. Note the form of the sensitivity which expresses the change in the logarithm of the surface temperature with the logarithm of the spectral emissivity is equivalent to the percentage change in the true surface temperature with the percentage change in the spectral emissivity. Quantifying the uncertainty in this manner simplifies the calculation of relative error.

## Slide 23: Spectral Method Error Estimation - III
Chart: True Surface Temperature T, K
| 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1100 | 1200 | 1300 | 1400 | 1500 | 1600 | 1700 | 1800 | 1900 | 2000 | 2100 | 2200 | 2300 | 2400 | 2500 | 2600 | 2700 | 2800 | 2900 | 3000 | 3100 | 3200 | 3300 | 3400 | 3500 | 3600 | 3700 | 3800 | 3900 | 4000 | 4100 | 4200 | 4300 | 4400 | 4500 | 4600 | 4700 | 4800 | 4900 | 5000
0.5 | -0.010426 | -0.013901 | -0.017376 | -0.020851 | -0.024326 | -0.027801 | -0.031277 | -0.034752 | -0.038227 | -0.041702 | -0.045177 | -0.048652 | -0.052128 | -0.055603 | -0.059078 | -0.062553 | -0.066028 | -0.069504 | -0.072979 | -0.076454 | -0.079929 | -0.083404 | -0.086879 | -0.090353 | -0.093828 | -0.097302 | -0.10078 | -0.10425 | -0.10772 | -0.11119 | -0.11466 | -0.11813 | -0.1216 | -0.12506 | -0.12853 | -0.13199 | -0.13545 | -0.1389 | -0.14235 | -0.1458 | -0.14925 | -0.15269 | -0.15612 | -0.15955 | -0.16298 | -0.16639 | -0.1698 | -0.17321
1 | -0.020851 | -0.027801 | -0.034752 | -0.041702 | -0.048652 | -0.055603 | -0.062553 | -0.069504 | -0.076454 | -0.083404 | -0.090353 | -0.097302 | -0.10425 | -0.11119 | -0.11813 | -0.12506 | -0.13199 | -0.1389 | -0.1458 | -0.15269 | -0.15955 | -0.16639 | -0.17321 | -0.18 | -0.18675 | -0.19347 | -0.20015 | -0.20679 | -0.21338 | -0.21993 | -0.22643 | -0.23288 | -0.23927 | -0.24561 | -0.2519 | -0.25812 | -0.26429 | -0.27039 | -0.27644 | -0.28242 | -0.28834 | -0.29419 | -0.29998 | -0.30571 | -0.31137 | -0.31696 | -0.3225 | -0.32796
2 | -0.041702 | -0.055603 | -0.069504 | -0.083404 | -0.097302 | -0.11119 | -0.12506 | -0.1389 | -0.15269 | -0.16639 | -0.18 | -0.19347 | -0.20679 | -0.21993 | -0.23288 | -0.24561 | -0.25812 | -0.27039 | -0.28242 | -0.29419 | -0.30571 | -0.31696 | -0.32796 | -0.3387 | -0.34918 | -0.35941 | -0.36938 | -0.37911 | -0.3886 | -0.39785 | -0.40687 | -0.41566 | -0.42423 | -0.43259 | -0.44073 | -0.44868 | -0.45642 | -0.46398 | -0.47135 | -0.47853 | -0.48555 | -0.49239 | -0.49907 | -0.50559 | -0.51195 | -0.51816 | -0.52423 | -0.53016
4 | -0.083404 | -0.11119 | -0.1389 | -0.16639 | -0.19347 | -0.21993 | -0.24561 | -0.27039 | -0.29419 | -0.31696 | -0.3387 | -0.35941 | -0.37911 | -0.39785 | -0.41566 | -0.43259 | -0.44868 | -0.46398 | -0.47853 | -0.49239 | -0.50559 | -0.51816 | -0.53016 | -0.54161 | -0.55255 | -0.563 | -0.573 | -0.58258 | -0.59175 | -0.60054 | -0.60898 | -0.61708 | -0.62486 | -0.63235 | -0.63954 | -0.64647 | -0.65315 | -0.65958 | -0.66579 | -0.67178 | -0.67756 | -0.68314 | -0.68854 | -0.69376 | -0.69882 | -0.70371 | -0.70845 | -0.71304
8 | -0.16639 | -0.21993 | -0.27039 | -0.31696 | -0.35941 | -0.39785 | -0.43259 | -0.46398 | -0.49239 | -0.51816 | -0.54161 | -0.563 | -0.58258 | -0.60054 | -0.61708 | -0.63235 | -0.64647 | -0.65958 | -0.67178 | -0.68314 | -0.69376 | -0.70371 | -0.71304 | -0.7218 | -0.73006 | -0.73784 | -0.7452 | -0.75215 | -0.75874 | -0.765 | -0.77094 | -0.77658 | -0.78196 | -0.78709 | -0.79199 | -0.79666 | -0.80113 | -0.80541 | -0.80951 | -0.81345 | -0.81722 | -0.82085 | -0.82433 | -0.82768 | -0.83091 | -0.83402 | -0.83702 | -0.83991
 | -0.010426 | -0.013901 | -0.017376 | -0.020851 | -0.024326 | -0.027801 | -0.031277 | -0.034752 | -0.038227 | -0.041702 | -0.045177 | -0.048652 | -0.052128 | -0.055603 | -0.059078 | -0.062553 | -0.066028 | -0.069504 | -0.072979 | -0.076454 | -0.079929 | -0.083404 | -0.086879 | -0.090355 | -0.09383 | -0.097305 | -0.10078 | -0.10426 | -0.10773 | -0.11121 | -0.11468 | -0.11816 | -0.12163 | -0.12511 | -0.12858 | -0.13206 | -0.13553 | -0.13901 | -0.14248 | -0.14596 | -0.14943 | -0.15291 | -0.15638 | -0.15986 | -0.16333 | -0.16681 | -0.17028 | -0.17376
 | -0.020851 | -0.027801 | -0.034752 | -0.041702 | -0.048652 | -0.055603 | -0.062553 | -0.069504 | -0.076454 | -0.083404 | -0.090355 | -0.097305 | -0.10426 | -0.11121 | -0.11816 | -0.12511 | -0.13206 | -0.13901 | -0.14596 | -0.15291 | -0.15986 | -0.16681 | -0.17376 | -0.18071 | -0.18766 | -0.19461 | -0.20156 | -0.20851 | -0.21546 | -0.22241 | -0.22936 | -0.23631 | -0.24326 | -0.25021 | -0.25716 | -0.26411 | -0.27106 | -0.27801 | -0.28496 | -0.29191 | -0.29887 | -0.30582 | -0.31277 | -0.31972 | -0.32667 | -0.33362 | -0.34057 | -0.34752
 | -0.041702 | -0.055603 | -0.069504 | -0.083404 | -0.097305 | -0.11121 | -0.12511 | -0.13901 | -0.15291 | -0.16681 | -0.18071 | -0.19461 | -0.20851 | -0.22241 | -0.23631 | -0.25021 | -0.26411 | -0.27801 | -0.29191 | -0.30582 | -0.31972 | -0.33362 | -0.34752 | -0.36142 | -0.37532 | -0.38922 | -0.40312 | -0.41702 | -0.43092 | -0.44482 | -0.45872 | -0.47262 | -0.48652 | -0.50043 | -0.51433 | -0.52823 | -0.54213 | -0.55603 | -0.56993 | -0.58383 | -0.59773 | -0.61163 | -0.62553 | -0.63943 | -0.65333 | -0.66723 | -0.68113 | -0.69504
 | -0.083404 | -0.11121 | -0.13901 | -0.16681 | -0.19461 | -0.22241 | -0.25021 | -0.27801 | -0.30582 | -0.33362 | -0.36142 | -0.38922 | -0.41702 | -0.44482 | -0.47262 | -0.50043 | -0.52823 | -0.55603 | -0.58383 | -0.61163 | -0.63943 | -0.66723 | -0.69504 | -0.72284 | -0.75064 | -0.77844 | -0.80624 | -0.83404 | -0.86184 | -0.88965 | -0.91745 | -0.94525 | -0.97305 | -1.0009 | -1.0287 | -1.0565 | -1.0843 | -1.1121 | -1.1399 | -1.1677 | -1.1955 | -1.2233 | -1.2511 | -1.2789 | -1.3067 | -1.3345 | -1.3623 | -1.3901
 | -0.16681 | -0.22241 | -0.27801 | -0.33362 | -0.38922 | -0.44482 | -0.50043 | -0.55603 | -0.61163 | -0.66723 | -0.72284 | -0.77844 | -0.83404 | -0.88965 | -0.94525 | -1.0009 | -1.0565 | -1.1121 | -1.1677 | -1.2233 | -1.2789 | -1.3345 | -1.3901 | -1.4457 | -1.5013 | -1.5569 | -1.6125 | -1.6681 | -1.7237 | -1.7793 | -1.8349 | -1.8905 | -1.9461 | -2.0017 | -2.0573 | -2.1129 | -2.1685 | -2.2241 | -2.2797 | -2.3353 | -2.3909 | -2.4465 | -2.5021 | -2.5577 | -2.6133 | -2.6689 | -2.7245 | -2.7801
Notes: Here we show the sensitivity of the measured true surface temperature to the spectral emissivity as a function of the true surface temperature with wavelength as a parameter. For comparison, the sensitivity calculated using Wien’s approximation is also shown. Note that the magnitude of the sensitivity calculated using Wien’s approximation is always greater than the value derived from the full Planck Equation. Thus, Wien’s approximation provides an upper bound on the magnitude of the sensitivity.

## Slide 24: Example 6 - I
E6: If the spectral emissivity of graphite at 0.53 m is estimated to be 0.8 with an estimated uncertainty of 20%, what is the estimated uncertainty in the temperature if a detector at this wavelength indicates an equivalent blackbody temperature of 2,950 K?
Notes: Example 6. If the spectral emissivity of graphite at 0.53 m is estimated to be 0.8 with an estimated uncertainty of 20%, what is the estimated uncertainty in the temperature if a detector at this wavelength indicates an equivalent blackbody temperature of 2,950 K?
Example 6. If the spectral emissivity of graphite at 0.53 𝜇m is estimated to be 0.8 with an estimated uncertainty of 20%, what is the estimated uncertainty in the temperature if a detector at this wavelength indicates an equivalent blackbody temperature of 2,950 K?

## Slide 25: Example 6 - II
A6: The wavelength temperature product is 2,950 K times 0.53 m or 1,563.5 m-K. This is much less than or 14,388 m-K. So, Wien’s approximation is valid and Equation 14 and routine bb_tstw can be used:
Therefore, since the emissivity uncertainty is estimated to be 20%, the resulting uncertainty in surface temperature is 20% × 11% or 2.2%. For a temperature of 3,023 K, the absolute uncertainty is 67 K.
Notes: The wavelength temperature product is 2,950 K times 0.53 m or 1,563.5 m-K. This is much less than or 14,388 m-K. So, Wien’s approximation is valid and Equation 14 and routine bb_tstw can be used. The calculated true surface temperature is 3,023 K and the sensitivity to the spectral emissivity is 11%. Therefore, since the emissivity uncertainty is estimated to be 20%, the resulting uncertainty in surface temperature is 20% × 11% or 2.2%. For a temperature of 3,023 K, the absolute uncertainty is 67 K.
The wavelength temperature product is 2,950 K times 0.53 𝜇m or 1,563.5 𝜇m-K. This is much less than 𝐶_2 or 14,388 m-K. So, Wien’s approximation is valid and Equation 14 and routine bb_tstw can be used. The calculated true surface temperature is 3,023 K and the sensitivity to the spectral emissivity is 11%.
Therefore, since the emissivity uncertainty is estimated to be 20%, the resulting uncertainty in surface temperature is 20% × 11% or 2.2%. For a temperature of 3,023 K, the absolute uncertainty is 67 K.

## Slide 26: Example 7 - I
E7: Repeat example 6 except for a detector with a wavelength of 5.8 m where the emissivity is estimated to be 0.8 with an estimated uncertainty of 20% at an equivalent blackbody temperature of 2,950 K.
Notes: Example 7. Repeat example 6 except for a detector with a wavelength of 5.8 m where the emissivity is estimated to be 0.8 with an estimated uncertainty of 20% at an equivalent blackbody a temperature of 2,950 K.
Example 7. Repeat example 6 except for a detector with a wavelength of 5.8 𝜇m where the emissivity is estimated to be 0.8 with an estimated uncertainty of 20% at a temperature of 2,950 K.

## Slide 27: Example 7 - II
A7: The wavelength temperature product is 2,950 K times 5.8 m or 17,110 m-K. This is greater than or 14,388 m-K, so, Wien’s approximation is not valid and the full Planck equation (Equation 1) and routine bb_tst must be used.
Therefore, since the emissivity uncertainty is estimated to be 20%, the resulting uncertainty in temperature is 20% × 71% or 14%. For a temperature of 3,445 K, the absolute uncertainty is about 491 K!
Notes: The wavelength temperature product is 2,950 K times 5.8 m or 17,110 m-K. This is greater than or 14,388 m-K, so, Wien’s approximation is not valid and the full Planck equation (Equation 1) and routine bb_tst must be used. The calculated true surface temperature is 3,445 K and the sensitivity to the spectral emissivity is 71%. Therefore, since the emissivity uncertainty is estimated to be 20%, the resulting uncertainty in temperature is 20% × 71% or 14%. For a temperature of 3,445 K, the absolute uncertainty is about 491 K!
The wavelength temperature product is 2,950 K times 5.8 𝜇m or 17,110 𝜇m-K. This is greater than 𝐶_2 or 14,388 m-K, so, Wien’s approximation is not valid and the full Planck equation (Equation 1) and routine bb_tst must be used. The calculated true surface temperature is 3,445 K and the sensitivity to the spectral emissivity is 79%.
Therefore, since the emissivity uncertainty is estimated to be 20%, the resulting uncertainty in temperature is 20% × 79% or 16%. For a temperature of 3,445 K, the absolute uncertainty is about 545 K!

## Slide 28: Rule 1
When the emissivity is unknown and must be estimated, the most accurate surface temperature measurement is made when a detector with a wavelength as short as possible is used.
Notes: The previous analysis and examples lead us to the first of four rules for radiation thermometry. Rule 1: When the emissivity is unknown and must be estimated, the most accurate surface temperature measurement is made when a detector with a wavelength as short as possible is used.

## Slide 29: Determination of Emissivities at Other Wavelengths
Once the true surface temperature has been determined from a detector at one wavelength, the emissivity at other detector wavelengths can be determined from:
Of course, using a measurement at the same wavelength will return our initial guess for .
 |  | (18)
 |  | (18)
Notes: Once the true surface temperature has been determined from a detector at one wavelength, the emissivity at other detector wavelengths can be determined from the equation shown. Of course, using a measurement at the same wavelength will return our initial guess for .
Once the true surface temperature has been determined from a detector at one wavelength, the emissivity at other detector wavelengths can be determined from the equation shown. Of course, using a measurement at the same wavelength will return our initial guess for 𝜀_𝜆.

## Slide 30: Emissivity Uncertainty Estimation - I
The uncertainty in the emissivity at a second wavelength based on the estimated uncertainty in a given wavelength is, by differentiation:
If Wien’s approximation is valid at both wavelengths then:
 |  | (19)
 |  | (19)
 |  | (20)
 |  | (20)
Notes: The uncertainty in the emissivity at a second wavelength based on the estimated uncertainty in the emissivity at a given wavelength is, determined by differentiation and shown in Equation 19. If Wien’s approximation is valid at both wavelengths then the sensitivity becomes proportional to the ratio of the first wavelength to the second wavelength (Equation 20).
The uncertainty in the emissivity at a second wavelength 𝜀_(𝜆_𝑖 )based on the estimated uncertainty in a given wavelength 𝜀_(𝜆_1 )is, by differentiation shown in Equation 19. If Wien’s approximation is valid at both wavelengths then the sensitivity becomes proportional to the ratio of the first wavelength to the second wavelength.

## Slide 31: Emissivity Uncertainty Estimation - II
Unlike with temperature, the uncertainty in the computed emissivity decreases with increasing wavelength.
The error is approximately inversely proportional to the ratio of the two wavelengths.
Therefore emissivities derived from longer wavelength detectors have less uncertainty. However, since the emissivity is usually a strong function of wavelength, values at long wavelengths may or may not be useful if emissivities over the entire wavelength band are needed.
Notes: Unlike temperature, the uncertainty in the computed emissivity at a second wavelength decreases with increasing wavelength. The error is approximately inversely proportional to the ratio of the two wavelengths. Therefore emissivities derived from longer wavelength detectors have less uncertainty. However, since the emissivity is usually a strong function of wavelength, values at long wavelengths may or may not be useful if emissivities over the entire wavelength band are needed.

## Slide 32: Emissivity Uncertainty Estimation - III
Chart: True Surface Temperature T, K
| 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1100 | 1200 | 1300 | 1400 | 1500 | 1600 | 1700 | 1800 | 1900 | 2000 | 2100 | 2200 | 2300 | 2400 | 2500 | 2600 | 2700 | 2800 | 2900 | 3000 | 3100 | 3200 | 3300 | 3400 | 3500 | 3600 | 3700 | 3800 | 3900 | 4000 | 4100 | 4200 | 4300 | 4400 | 4500 | 4600 | 4700 | 4800 | 4900 | 5000

 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.50001 | 0.50002 | 0.50003 | 0.50006 | 0.50011 | 0.50017 | 0.50026 | 0.50038 | 0.50053 | 0.50072 | 0.50096 | 0.50125 | 0.50158 | 0.50198 | 0.50242 | 0.50293 | 0.5035 | 0.50413 | 0.50482 | 0.50558 | 0.50639 | 0.50726 | 0.5082 | 0.50919 | 0.51024 | 0.51134 | 0.5125 | 0.5137 | 0.51496 | 0.51626 | 0.51761 | 0.519 | 0.52044 | 0.52191 | 0.52342 | 0.52496 | 0.52653 | 0.52814
 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25001 | 0.25003 | 0.25008 | 0.25019 | 0.25036 | 0.25062 | 0.25099 | 0.25148 | 0.25208 | 0.25282 | 0.25369 | 0.25468 | 0.2558 | 0.25704 | 0.25841 | 0.25988 | 0.26145 | 0.26313 | 0.2649 | 0.26676 | 0.26871 | 0.27073 | 0.27282 | 0.27498 | 0.2772 | 0.27948 | 0.28182 | 0.2842 | 0.28663 | 0.28911 | 0.29162 | 0.29417 | 0.29676 | 0.29937 | 0.30202 | 0.30469 | 0.30738 | 0.31009 | 0.31283 | 0.31558 | 0.31834 | 0.32112 | 0.32391 | 0.32671
 | 0.125 | 0.12502 | 0.12509 | 0.12531 | 0.12574 | 0.12641 | 0.12734 | 0.12852 | 0.12994 | 0.13157 | 0.13338 | 0.13537 | 0.1375 | 0.13976 | 0.14213 | 0.1446 | 0.14716 | 0.1498 | 0.1525 | 0.15527 | 0.15809 | 0.16096 | 0.16387 | 0.16682 | 0.16981 | 0.17283 | 0.17587 | 0.17894 | 0.18204 | 0.18515 | 0.18829 | 0.19144 | 0.1946 | 0.19778 | 0.20097 | 0.20417 | 0.20738 | 0.21059 | 0.21381 | 0.21704 | 0.22027 | 0.22351 | 0.22674 | 0.22998 | 0.23322 | 0.23645 | 0.23969 | 0.24292
 | 0.062656 | 0.063205 | 0.064261 | 0.065784 | 0.067684 | 0.069879 | 0.072302 | 0.0749 | 0.077636 | 0.080481 | 0.083413 | 0.086416 | 0.089478 | 0.092588 | 0.095738 | 0.098923 | 0.10214 | 0.10538 | 0.10864 | 0.11191 | 0.11521 | 0.11852 | 0.12184 | 0.12518 | 0.12852 | 0.13187 | 0.13523 | 0.1386 | 0.14197 | 0.14535 | 0.14873 | 0.15212 | 0.1555 | 0.15889 | 0.16229 | 0.16568 | 0.16907 | 0.17246 | 0.17585 | 0.17924 | 0.18263 | 0.18601 | 0.18939 | 0.19277 | 0.19614 | 0.19951 | 0.20287 | 0.20622


 | 0.5 | 0.5 | 0.5 | 0.5 | 0.50002 | 0.50006 | 0.50017 | 0.50038 | 0.50072 | 0.50125 | 0.50198 | 0.50293 | 0.50413 | 0.50558 | 0.50726 | 0.50919 | 0.51134 | 0.5137 | 0.51626 | 0.519 | 0.52191 | 0.52496 | 0.52814 | 0.53143 | 0.53482 | 0.5383 | 0.54184 | 0.54545 | 0.54911 | 0.5528 | 0.55652 | 0.56027 | 0.56402 | 0.56778 | 0.57154 | 0.5753 | 0.57905 | 0.58278 | 0.58649 | 0.59018 | 0.59384 | 0.59748 | 0.60109 | 0.60466 | 0.6082 | 0.61171 | 0.61518 | 0.61861
 | 0.25 | 0.25003 | 0.25019 | 0.25062 | 0.25148 | 0.25282 | 0.25468 | 0.25704 | 0.25988 | 0.26313 | 0.26676 | 0.27073 | 0.27498 | 0.27948 | 0.2842 | 0.28911 | 0.29417 | 0.29937 | 0.30469 | 0.31009 | 0.31558 | 0.32112 | 0.32671 | 0.33233 | 0.33798 | 0.34364 | 0.3493 | 0.35495 | 0.3606 | 0.36622 | 0.37182 | 0.37739 | 0.38292 | 0.38842 | 0.39387 | 0.39928 | 0.40464 | 0.40995 | 0.41521 | 0.42041 | 0.42555 | 0.43065 | 0.43568 | 0.44065 | 0.44557 | 0.45042 | 0.45522 | 0.45995
 | 0.12531 | 0.12641 | 0.12852 | 0.13157 | 0.13537 | 0.13976 | 0.1446 | 0.1498 | 0.15527 | 0.16096 | 0.16682 | 0.17283 | 0.17894 | 0.18515 | 0.19144 | 0.19778 | 0.20417 | 0.21059 | 0.21704 | 0.22351 | 0.22998 | 0.23645 | 0.24292 | 0.24937 | 0.2558 | 0.26221 | 0.26858 | 0.27493 | 0.28123 | 0.28749 | 0.29371 | 0.29988 | 0.30599 | 0.31205 | 0.31806 | 0.32401 | 0.32989 | 0.33572 | 0.34149 | 0.34719 | 0.35283 | 0.3584 | 0.36391 | 0.36935 | 0.37473 | 0.38004 | 0.38529 | 0.39047


Notes: This plot illustrates the sensitivity of derived emissivity for various detector combinations as a function of true surface temperature. The sensitivity at low temperature approaches the value calculated from Wien’s approximation however at high temperatures, the sensitivity calculated using Wien’s approximation is always less than the value derived from the full Planck Equation. Thus, Wien’s approximation provides a lower bound on the emissivity sensitivity.

## Slide 33: Example 8 - I
E8: A 0.53-m wavelength detector is used to determine the surface temperature. The measured equivalent blackbody temperature from this detector is 1,950 K and the emissivity is estimated to be 0.6 with an estimated uncertainty of 25%. What are the inferred emissivities and uncertainties at 1.0 m and 5.8 m if detectors at both wavelengths indicate equivalent blackbody temperatures of 1,740 K?
Notes: Example 8: A 0.53-m wavelength detector is used to determine the surface temperature. The measured equivalent blackbody temperature from this detector is 1,950 K and the emissivity is estimated to be 0.6 with an estimated uncertainty of 25%. What are the inferred emissivities and uncertainties at 1.0 m and 5.8 m if detectors at both wavelengths indicate equivalent blackbody temperatures of 1,740 K?

## Slide 34: Example 8 - II
A8: The true surface temperature using Equation 13 and routine bb_tst is:
The emissivity at 1.0 m using Equation 15 and routine bb_emiss is:
Notes: The true surface temperature using Equation 13 and routine bb_tst is equal to 2,024 K. The emissivity at 1.0 m using Equation 15 and routine bb_emiss is 0.313.

## Slide 35: Example 8 - III
The sensitivity to the emissivity at wavelength 2 based on the assumed emissivity at wavelength 1 using Equation 19 and routine bb_dlne2dlne1 is:
And the relative uncertainty in the emissivity is:
The emissivity at 5.8 m using routine bb_emiss is:
Notes: The sensitivity to the emissivity at wavelength 2 based on the assumed emissivity at wavelength 1 using Equation 19 and routine bb_dlne2dlne1 is calculated to be 0.530. The relative error in the emissivity is 13.3%. The emissivity at 5.8 m using routine bb_emiss is calculated to be 0.761.

## Slide 36: Example 8 - IV
The sensitivity to the emissivity at wavelength 2 based on the assumed emissivity at wavelength 1 using routine bb_dlne2dlne1 is:
And the relative uncertainty in the emissivity is:
Note that the emissivity uncertainty at 5.8 m is four times less than at 1m.
Notes: The sensitivity to the emissivity at wavelength 2 based on the assumed emissivity at wavelength 1 using routine bb_dlne2dlne1 is 0.129 and the relative uncertainty in the emissivity at 5.8 m is therefore 3.2%. Note that the emissivity uncertainty at 5.8 m is four times less than at 1 m.
The sensitivity to the emissivity at wavelength 2 based on the assumed emissivity at wavelength 1 using routine bb_dlne2dlne1 is 0.129 and the relative uncertainty in the emissivity at 5.8 𝜇m is therefore 3.2%. Note that the emissivity uncertainty at 5.8 𝜇m is four times less than at 1𝜇m.

## Slide 37: Rule 2
The uncertainty in the spectral emissivity at a given wavelength decreases approximately inversely with the wavelength that is used.
Notes: Unlike temperature, the uncertainty in emissivities at longer wavelengths is lower than at shorter wavelengths. This leads us to Rule 2: The uncertainty in the spectral emissivity at a given wavelength decreases approximately inversely with the wavelength that is used.

## Slide 38: Method 2 – Ratio method - I
Make measurements and with detectors at two different wavelengths and . Assume a ratio for the spectral emissivities at the two wavelengths (could be 1) and then calculate the true temperature by solving the following:
and then calculate the individual emissivities from:
 |  | (21)
 |  | (21)
 |  | (22)
 |  | (22)
Notes: The second method we will investigate is the ratio method. For this method, we make measurements and with detectors at two different wavelengths and . We then assume a ratio for the spectral emissivities at the two wavelengths (this could be 1). We then calculate the true temperature by solving for the emissivity ratio in terms of the ratio of the measured and assumed spectral intensities as shown in Equation 21. The individual emissivities can then be determined using Equation 22 as shown.
The second method we will investigate is the ratio method. For this method, we make measurements 𝑇_(𝜆_1 )and 𝑇_(𝜆_2 )at two different wavelengths 𝜆_1 and 𝜆_2. Assume a ratio for the spectral emissivities at the two wavelengths ε_𝑟(could be 1) and then calculate the true temperature 𝑇 by solving for the emissivity ratio in terms of the ratio of the measured and assumed spectral intensities. The individual emissivities can then be determined using Equation 22 as shown.

## Slide 39: Method 2 – Ratio method - II
If Wien’s approximation is valid, the approximate solution for the true temperature can be calculated directly by:
where is an equivalent wavelength
and is the ratio temperature, the temperature of an equivalent blackbody having the same ratio of spectral intensities at the two specified wavelengths as that of the target.
 |  | (23)
 |  | (23)
 |  | (24)
 |  | (24)
 |  | (25)
 |  | (25)
Notes: If Wien’s approximation is valid, the approximate solution for the true temperature can be calculated directly by Equation 23 as shown. In Equation 23, is an equivalent wavelength and is the ratio temperature, the temperature of an equivalent blackbody having the same ratio of spectral intensities at the two specified wavelengths as that of the target.
If Wien’s approximation is valid, the approximate solution for the true temperature 𝑇 can be calculated directly by Equation 22 as shown. In Equation 22, Λ is an equivalent wavelength and 𝑇_𝑟 is the ratio temperature, the temperature of an equivalent blackbody having the same ratio of spectral radiances at the two specified wavelengths as that of the target.

## Slide 40: Method 2 – Ratio method - III
Once the true surface temperature is known, the emissivities can then be determined by:
 |  | (26)
 |  | (26)
Notes: Once the true surface temperature is known, the emissivities can then be determined by Equation 26.

## Slide 41: Ratio Method Observations
For an emissivity ratio of 1, the calculated ratio temperature of the instrument is the true surface temperature.
The equivalent wavelength can be much greater than either of the two individual wavelengths when the difference between the two wavelengths is small.
As the difference between the two wavelengths increases, the value of the equivalent wavelength approaches the value of wavelength 1.
For Wien’s approximation to be valid, the wavelength temperature product of the two individual detectors need to be small, not the product of the equivalent temperature and the effective wavelength.
Equation 23 for the true surface temperature is identical to the equation for surface temperature in the spectral method (Equation 14) with the effective wavelength replacing the single wavelength and the ratio temperature replacing the single detector temperature.
Notes: The following observations can be drawn from the previous analysis:
1) For an emissivity ratio of 1, the calculated ratio temperature of the instrument is the true surface temperature.
2) The equivalent wavelength can be much greater than either of the two individual wavelengths when the difference between the two wavelengths is small. As the difference between the two wavelengths increases, the value of the equivalent wavelength approaches the value of wavelength 1.
3) For Wien’s approximation to be valid, the wavelength temperature product of the two individual detectors both need to be small, not just the product of the equivalent temperature and the effective wavelength.
4) Equation 23 for the true surface temperature is identical to Equation 14 for the surface temperature in the spectral method with two key differences, however. These are the effective wavelength in the ratio method replaces the single wavelength in the spectral method and the ratio temperature replaces the single detector temperature.

## Slide 42: Example 9 - I
E9: For a two-band radiation thermometer with detector wavelengths at 0.5 and 0.6 m, what are the true surface temperature and the spectral emissivities when the measured equivalent blackbody temperatures at the two wavelengths are 2,800 and 2,750 K, respectively? Assume an emissivity ratio of 0.9.
Notes: Example 9 here will demonstrate the ratio method. For a two-band radiation thermometer with detector wavelengths at 0.5 and 0.6 m, what is the true surface temperature and the spectral emissivities when the measured equivalent blackbody temperatures at the two wavelengths are 2,800 and 2,750 K, respectively? Assume an emissivity ratio of 0.9.

## Slide 43: Example 9 - II
A9: In this case, Wien’s approximation is valid so that using the approximate method from Equation 23 and calculating the effective wavelength from Equation 24 using routine bb_elam is:
Using Equation 25, the ratio temperature from routine bb_etemp is:
and using Equation 23 and routine bb_tstw, the true surface temperature is:
Notes: In this case, Wien’s approximation is valid so that using the approximate method from Equation 23 is acceptable. Calculating the effective wavelength from Equation 24 using routine bb_elam results in a value of 3 . Using Equation 25, the ratio temperature from routine bb_etemp is . Using Equation 23 and routine bb_tstw, the true surface temperature is then found to be 3,304 K.
In this case, Wien’s approximation is valid so that using the approximate method from Equation 22 and calculating the effective wavelength Λ from Equation 23 using routine bb_elam is 3 m. Using Equation 24, the ratio temperature 𝑇_𝑟 from routine bb_etemp is "3,080 K". Using Equation 22 and routine bb_tstw, the true surface temperature is 3,304 K.

## Slide 44: Example 9 - III
Knowing the individual measured detector temperatures and the surface temperature, the emissivities are determined using Equation 18 and routine bb_emiss:
So that, consistent with our initial assumption.
Notes: Knowing the individual measured detector temperatures and the calculated true surface temperature, the emissivities at 0.5 and 0.6 µm are determined using Equation 22 and routine bb_emiss which are 0.209 and 0.232, respectively.

## Slide 45: Ratio Method Sensitivities
The sensitivity of the true surface temperature to the emissivity ratio can be calculated for the case where Planck’s law is valid by differentiating Equation 17. This is:
When Wien’s approximation is valid, then the sensitivity can be easily determined utilizing the similarity to the spectral method as discussed before. The sensitivity then becomes:
 |  | (28)
 |  | (28)
 |  | (27)
 |  | (27)
Notes: Similar to the spectral method, the sensitivity of the true surface temperature to the emissivity ratio can be calculated for the case where Planck’s law is valid by differentiating Equation 17. This is shown in Equation 27. When Wien’s approximation is valid, then the sensitivity can be easily determined utilizing the similarity to the spectral method as discussed before and is shown in Equation 28.

## Slide 46: Example 10 - I
E10: For the same conditions as Example 9 (measured equivalent blackbody temperatures of 2,800 and 2,750 K for 0.5 and 0.6 m detectors, respectively), what would the error in the true surface temperature be if an assumed emissivity ratio of 1.0 was used but the actual ratio of the surface emissivities was 0.9?
Notes: Example 10. For the same conditions as Example 9, what would the error in the true surface temperature be if an assumed emissivity ratio of 1.0 was used but the actual ratio of the surface emissivities was 0.9?

## Slide 47: Example 10 - II
A10: This problem can be solved two ways: directly or using sensitivities. We will use sensitivities first.
In this case Wien’s approximation is valid since both the wavelength-temperature products are smaller than. The sensitivity from Equation 26 is:
So that the estimated uncertainty in temperature is:
Notes: This problem can be solved two ways: directly or using sensitivities. We will use sensitivities first. In this case Wien’s approximation is valid since both the wavelength-temperature products are smaller than. The sensitivity from Equation 26 is calculated to be 0.688. The estimated uncertainty in temperature is then 6.9% or 228 K.
This problem can be solved two ways: directly or using sensitivities. We will use sensitivities first.
In this case Wien’s approximation is valid since the both the wavelength-temperature products are small than 𝐶_2. The sensitivity from Equation 25 is calculated to be 0.688. The estimated uncertainty in temperature is then 6.9% or 228 K.

## Slide 48: Example 10 - III
Rather than using sensitivities, we can calculate the difference directly:
For an emissivity of 1.0, the true surface temperature is just the ratio temperature, or 3,080 K. Therefore, the error is:
almost exactly equal to the linear approximation using sensitivities.
Notes: Rather than using sensitivities, we can calculate the difference directly. For an emissivity of 1.0, the true surface temperature is just the ratio temperature, or 3080 K. Therefore, the error is 224K. And almost exactly equal to the linear approximation using sensitivities.

## Slide 49: Example 11 - I
E11: For a two-band radiation thermometer with detector wavelengths at 4.0 and 8.0 m, what is the error introduced when an emissivity ratio of 1.0 is assumed, when in fact the emissivity ratio of the sample has a ratio of 0.9? Assume the measured spectral temperatures at the two wavelengths are 2,800 and 2,750 K, the same as in Problem 9.
Notes: Example 11. For a two-band radiation thermometer with detector wavelengths at 4.0 and 8.0 m, what is the error introduced when an emissivity ratio of 1.0 is assumed, when in fact the emissivity ratio of the sample is 0.9? Assume the measured spectral temperatures at the two wavelengths are 2,800 and 2,750 K, the same as in Problem 9.

## Slide 50: Example 11 - II
A11: Since the products of and for both wavelengths are not less than or , m-K, the full Planck Equation must be used to solve for the true surface temperature (Equation 31 and routine bb_tratio).
Notes: In this case, the effective wavelength from Equation 24 and using routine bb_elam is Since the products of and for both wavelengths are not less than or , m-K, the full Planck Equation must be used as shown above.
In this case, the effective wavelength Λ Equation 35 and using routine bb_elam is 8.0 m.
Since the products of 𝑇_𝜆 and 𝜆 for both wavelengths are not less than 𝐶_(2 )or 14,388, 𝜇m-K, the full Planck Equation must be used to solve for the true surface temperature (Equation 21 and routine bb_tratio).

## Slide 51: Example 11 - III
The calculated values for true temperatures are:
The difference is 1,144 K!
This example points out the large error using long wavelengths and the need to use short wavelengths whenever possible.
Notes: The calculated values for true temperatures are 4,118 K for an assumed emissivity ratio of 0.9 and 2,974 K for an assumed emissivity ratio of 1.0. The difference is 1,144 K! This example points out the large error using long wavelengths and the need to use short wavelengths whenever possible.

## Slide 52: Example 11 - IV
The sensitivity of the emissivity ratio to surface temperature is (Equation 27):
Since the temperature change is so great, linear sensitivities extrapolated from one point will not be accurate.
However, the calculated sensitivities are:
Notes: The sensitivity of the emissivity ratio to surface temperature is given by Equation 27 and results in calculated sensitivities of -3.77 and -2.56 for the two wavelengths. However, since the temperature change is so great, linear sensitivities extrapolated from one point will not be accurate.

## Slide 53: Example 11 - V
The temperature error can be estimated using a centered average of the temperature and sensitivity:
almost exactly the same as the direct method.
Notes: So the temperature error can be estimated using a centered average of the temperature and sensitivity. The estimated uncertainty is 1,122 K almost exactly the same as the direct method.

## Slide 54: Further Observations
The selection of the optimal detector wavelengths for the ratio method is based upon two competing factors. As difference in the two detector wavelengths approaches zero the value of the effective wavelength goes to infinity and therefore tends to drive the temperature uncertainty higher. However, at the same time the smaller the difference in wavelengths makes the estimation of the emissivity ratio more accurate and therefore drives the temperature uncertainty lower.
Notes: Based on the previous analysis and the examples presented, we can see that the selection of the optimal detector wavelengths for the ratio method is based upon two competing factors. As the difference in the two detector wavelengths approaches zero, the value of the effective wavelength goes to infinity and therefore tends to drive the temperature uncertainty higher. However, at the same time the smaller the difference in wavelengths makes the estimation of the emissivity ratio more accurate and therefore drives the temperature uncertainty lower.

## Slide 55: Rule 3
The optimal detector wavelengths for the ratio method are based on two competing factors: 1) the need to keep the wavelength difference small to ensure an accurate assumed ratio, but 2) not too small that the effective wavelength increases and becomes too large.
Notes: Rule 3. The optimal detector wavelengths for the ratio method are based on two competing factors: 1) the need to keep the wavelength difference small to ensure an accurate assumed ratio, but 2) not too small that the effective wavelength increases and becomes too large.

## Slide 56: Multispectral Methods - I
Multispectral methods are an extension of the ratio method where a larger number of detectors are used.
Multiple measurements are then used to perform a best fit to a single temperature and wavelength relationship.
The emissivity relationship can be a constant value or some other more complex function such as a polynomial.
Notes: The final technique we will discuss are multispectral methods. Multispectral methods are an extension of the ratio method where a larger number of detectors are used. Multiple measurements are then used to perform a best fit to a single temperature and wavelength relationship. The emissivity relationship can be a constant value or some other more complex function such as a polynomial.

## Slide 57: Multispectral Methods - II
Advantages
  Large number of detectors reduces noise and averages errors inherent in measurements
Disadvantages
  Requires more complex hardware
  Multiple detectors increases data collection requirements and requires increased processing
  Data may not provide an identifiably unique solution unless a large number of measurements are made
Notes: The primary advantage of multispectral methods is that a large number of detectors reduces noise and averages errors inherent in measurements. The disadvantages are: 1) it requires more complex hardware, 2) multiple detectors increases data collection requirements and requires increased processing, and 3) the data may not provide an identifiably unique solution unless a large number of measurements are made. Given that in recent years the cost of detectors, electronics, and computer processing power have decreased dramatically, the potential advantage that multispectral methods bring usually far outweigh the disadvantages.

## Slide 58: Multispectral Hardware Utilizing Dispersive Spectrometer and Silicon Array Detector
Focusing Mirrors
Grating
Array Detector
Entrance Slit
32-Channel Array Silicon Detector
Notes: Here we present an example of a multispectral instrument using a dispersive spectrometer and an array detector. The grating separates the incoming light into the component wavelengths which are focused on the array. In this instrument, the detector has 32 separate channels each measuring a distinct wavelength. While this instrument only has provisions for 32 channels, it is possible to procure array detectors with over 1,000 elements and channels. The method outlined here can easily be extended to accommodate as many wavelengths as desired.

## Slide 59: Example 12 - I
E12: Given data to the right from a 32-channel silicon array detector, calculate the spectral emissivity and true surface temperature if the spectral emissivity is assumed to follow a second-order polynomial with wavelength.
Chart: Wavelength, m
| 0.32500000000000001 | 0.35000000000000003 | 0.37500000000000006 | 0.40000000000000008 | 0.4250000000000001 | 0.45000000000000012 | 0.47500000000000014 | 0.50000000000000011 | 0.52500000000000013 | 0.55000000000000016 | 0.57500000000000018 | 0.6000000000000002 | 0.62500000000000022 | 0.65000000000000024 | 0.67500000000000027 | 0.70000000000000029 | 0.72500000000000031 | 0.75000000000000033 | 0.77500000000000036 | 0.80000000000000038 | 0.8250000000000004 | 0.85000000000000042 | 0.87500000000000044 | 0.90000000000000047 | 0.92500000000000049 | 0.95000000000000051 | 0.97500000000000053 | 1.0000000000000004 | 1.0250000000000004 | 1.0500000000000003 | 1.0750000000000002 | 1.1000000000000001
Intensity | 0.21609 | 0.33496 | 0.50658 | 0.67537 | 0.87045 | 1.0946 | 1.2652 | 1.491 | 1.6397 | 1.8358 | 1.9845 | 2.0608 | 2.2122 | 2.2843 | 2.377 | 2.3759 | 2.4091 | 2.3908 | 2.4638 | 2.3899 | 2.3426 | 2.3857 | 2.3331 | 2.267 | 2.1678 | 2.1126 | 2.1016 | 2.0558 | 2.0018 | 1.9129 | 1.8322 | 1.7653
Eq BB Temperature | 3710.2847 | 3695.7916 | 3701.8937 | 3687.6531 | 3680.8611 | 3682.534 | 3663.863 | 3667.2086 | 3650.2149 | 3652.2946 | 3646.1675 | 3626.4823 | 3632.2314 | 3622.7106 | 3623.6845 | 3605.1103 | 3599.8962 | 3584.71 | 3598.8761 | 3574.7481 | 3560.5202 | 3577.1188 | 3566.4796 | 3552.8935 | 3528.3684 | 3520.2935 | 3530.8692 | 3529.4778 | 3525.6499 | 3506.5304 | 3490.2074 | 3479.8502
Notes: Example 12: Given data to the right from the 32-channel silicon array detector in the instrument just described, calculate the spectral emissivity and true surface temperature if the spectral emissivity is assumed to follow a second-order polynomial with wavelength.

## Slide 60: Example 12 - II
A12: The solution requires that we find values for ,, , and that minimize the function :
where
 |  | (29)
 |  | (29)
 |  | (30)
 |  | (30)
Notes: The solution requires that we find values for ,, , and that minimize the function which is the summation of the square difference between the measured and calculated spectral intensity at each of the 32 wavelengths. The measured spectral intensities are expressed in terms of the measured equivalent spectral blackbody temperatures. The emissivity is assumed to be expressed as a 2nd degree polynomial in wavelength.
The solution requires that we find values for 𝑇,〖 𝑎〗_0, 𝑎_1, and 𝑎_2 that minimize the function 𝑓 which is the summation of the square difference between the measured and calculated spectral intensity at each of the 32 wavelengths. The measured spectral intensities are expressed in terms of the measured equivalent spectral blackbody temperatures. The emissivity is assumed to be expressed as a polynomial in wavelength.

## Slide 61: Example 12 - III
Results:
Chart: Wavelength, m
| 0.32500000000000001 | 0.35000000000000003 | 0.37500000000000006 | 0.40000000000000008 | 0.4250000000000001 | 0.45000000000000012 | 0.47500000000000014 | 0.50000000000000011 | 0.52500000000000013 | 0.55000000000000016 | 0.57500000000000018 | 0.6000000000000002 | 0.62500000000000022 | 0.65000000000000024 | 0.67500000000000027 | 0.70000000000000029 | 0.72500000000000031 | 0.75000000000000033 | 0.77500000000000036 | 0.80000000000000038 | 0.8250000000000004 | 0.85000000000000042 | 0.87500000000000044 | 0.90000000000000047 | 0.92500000000000049 | 0.95000000000000051 | 0.97500000000000053 | 1.0000000000000004 | 1.0250000000000004 | 1.0500000000000003 | 1.0750000000000002 | 1.1000000000000001
Intensity Data | 0.21609 | 0.33496 | 0.50658 | 0.67537 | 0.87045 | 1.0946 | 1.2652 | 1.491 | 1.6397 | 1.8358 | 1.9845 | 2.0608 | 2.2122 | 2.2843 | 2.377 | 2.3759 | 2.4091 | 2.3908 | 2.4638 | 2.3899 | 2.3426 | 2.3857 | 2.3331 | 2.267 | 2.1678 | 2.1126 | 2.1016 | 2.0558 | 2.0018 | 1.9129 | 1.8322 | 1.7653
intensity Fit | 0.21563 | 0.34219 | 0.49857 | 0.67871 | 0.87476 | 1.0783 | 1.2813 | 1.4768 | 1.6594 | 1.825 | 1.971 | 2.0959 | 2.1994 | 2.2818 | 2.344 | 2.3875 | 2.4138 | 2.4246 | 2.4218 | 2.4069 | 2.3818 | 2.3479 | 2.3068 | 2.2595 | 2.2075 | 2.1516 | 2.0929 | 2.032 | 1.9697 | 1.9066 | 1.8432 | 1.7799
Emissivity Fit | 0.74795 | 0.74843 | 0.74883 | 0.74915 | 0.7494 | 0.74956 | 0.74965 | 0.74966 | 0.74959 | 0.74945 | 0.74923 | 0.74893 | 0.74855 | 0.74809 | 0.74756 | 0.74695 | 0.74626 | 0.74549 | 0.74465 | 0.74372 | 0.74272 | 0.74165 | 0.74049 | 0.73926 | 0.73794 | 0.73655 | 0.73509 | 0.73354 | 0.73192 | 0.73022 | 0.72844 | 0.72659
Notes: The solution is shown in the plot. The surface temperature of 3,802 K and the emissivity function provide a good fit to the data.

## Slide 62: Wide-Bandwidth Detectors
“The laws of light and of heat translate each other;—so do the laws of sound and colour; and so galvanism, electricity and magnetism are varied forms of this selfsame energy.” — Ralph Waldo Emerson
Notes: In this section, we will discuss wide-bandwidth detectors. Generally, these will involve measurements where the detector is used alone without a separate wavelength selection component.

## Slide 63: Wide-Bandwidth Detectors - I
Up until now, we have considered only narrow-band detectors, ones in which the wavelength band has been limited.
Sometime, using wide-band detectors prove to be an advantage.
However, using wide-band detectors is more difficult and requires additional understanding to properly interpret the measured data.
Notes: Up until now, we have considered only narrow-band detectors, ones in which the wavelength band has been eliminated.
Sometimes, using wide-band detectors proves to be an advantage. However, using wide-band detectors is more difficult and requires additional understanding to properly interpret the measured data

## Slide 64: Wide-Bandwidth Detectors - II
Advantages
  Provide a higher signal since the signal is measured over a wider band.
  Require less hardware since the wavelength limiting device is eliminated.
Disadvantages
  Introduces non-linearities in the calibration because the blackbody intensity varies with wavelength and the response of the detector is not uniform across all wavelengths.
  Makes data reduction and analysis more complex.
Notes:   Using wide-bandwidth detectors has several advantages. First, they provide a higher signal since the signal is measured over a wider bandwidth. Second, they require less hardware since the wavelength limiting device is eliminated. The primary disadvantage is that it introduces non-linearities in the calibration because the blackbody intensity varies with wavelength and the response of the detector is not uniform across all wavelengths. This makes data reduction and analysis more complex.

## Slide 65: Wide-Bandwidth Detectors - III
The spectral response curves as a function of wavelength are available for many detectors from manufacturers.
Often, though, these are only nominal specifications and can vary with the actual detector and can change over time for some detectors.
Measuring the response function requires more specialized equipment and methods not available in most laboratories, so that the exact response is often not available.
However, only estimates of spectral response are needed since many times the difference between the actual and nominal response is sufficient.
Notes: The spectral response curves as a function of wavelength are available for many detectors from manufacturers. Often, though, these are only nominal specifications and can vary with the actual detector and can change over time for some detectors. Measuring the response function requires more specialized equipment and methods not available in most laboratories, so that the exact response is often not available. However, only estimates of spectral responses are needed since many time the difference between the actual and nominal response is sufficient.
The spectral response curves as a function of wavelength 𝐷(𝜆) are available for many detectors from manufacturers. Often, though, these are only nominal specification and can vary with the actual detector and can change over time for some detectors.
Measuring the response function requires more specialized equipment and methods not available in most laboratories, so that the exact response is often not available. However, only estimates of spectral response is needed since many time the difference between the actual and nominal response is sufficient.

## Slide 66: Wide-Bandwidth Detectors - III
Non-linearities in the calibration are introduced because the measured signal is the net product of the detector response multiplied by the blackbody emission intensity.
Chart: 
| 0.33040000000000003 | 0.3478 | 0.36520000000000002 | 0.3826 | 0.40649999999999997 | 0.42170000000000002 | 0.43909999999999999 | 0.45429999999999998 | 0.46739999999999998 | 0.48480000000000001 | 0.5 | 0.51739999999999997 | 0.53700000000000003 | 0.55430000000000001 | 0.57169999999999999 | 0.59130000000000005 | 0.61519999999999997 | 0.63480000000000003 | 0.65429999999999999 | 0.67390000000000005 | 0.69350000000000001 | 0.73040000000000005 | 0.75 | 0.77170000000000005 | 0.7913 | 0.80430000000000001 | 0.82609999999999995 | 0.84570000000000001 | 0.86960000000000004 | 0.8891 | 0.90429999999999999 | 0.92610000000000003 | 0.94569999999999999 | 0.96740000000000004 | 0.98699999999999999 | 1 | 1.0086999999999999 | 1.0152000000000001 | 1.0196000000000001 | 1.0283 | 1.0326 | 1.0412999999999999 | 1.0478000000000001 | 1.0544 | 1.0587 | 1.0673999999999999 | 1.0761000000000001 | 1.087 | 1.0956999999999999 | 1.1022000000000001 | 1.1399999999999999
 | 0.039889 | 0.080124 | 0.12036 | 0.16059 | 0.21586 | 0.25101 | 0.29124 | 0.32639 | 0.35668 | 0.39691 | 0.43206 | 0.4723 | 0.51762 | 0.55762 | 0.5978 | 0.64267 | 0.69623 | 0.73872 | 0.77925 | 0.81783 | 0.85381 | 0.91304 | 0.93919 | 0.9632 | 0.97995 | 0.98826 | 0.99682 | 0.99833 | 0.99164 | 0.97872 | 0.96371 | 0.93423 | 0.89927 | 0.85071 | 0.79748 | 0.75706 | 0.72766 | 0.70444 | 0.6881 | 0.65431 | 0.63686 | 0.60004 | 0.57119 | 0.5407 | 0.52017 | 0.47705 | 0.43174 | 0.37185 | 0.32151 | 0.2824 | 0.028746
Chart: 
| 0.26519999999999999 | 0.28039999999999998 | 0.29780000000000001 | 0.31090000000000001 | 0.33040000000000003 | 0.3478 | 0.36520000000000002 | 0.3826 | 0.40649999999999997 | 0.42170000000000002 | 0.43909999999999999 | 0.45429999999999998 | 0.46739999999999998 | 0.48480000000000001 | 0.5 | 0.51739999999999997 | 0.53700000000000003 | 0.55430000000000001 | 0.57169999999999999 | 0.59130000000000005 | 0.61519999999999997 | 0.63480000000000003 | 0.65429999999999999 | 0.67390000000000005 | 0.69350000000000001 | 0.73040000000000005 | 0.75 | 0.77170000000000005 | 0.7913 | 0.80430000000000001 | 0.82609999999999995 | 0.84570000000000001 | 0.86960000000000004 | 0.8891 | 0.90429999999999999 | 0.92610000000000003 | 0.94569999999999999 | 0.96740000000000004 | 0.98699999999999999 | 1 | 1.0086999999999999 | 1.0152000000000001 | 1.0196000000000001 | 1.0283 | 1.0326 | 1.0412999999999999 | 1.0478000000000001 | 1.0544 | 1.0587 | 1.0673999999999999 | 1.0761000000000001 | 1.087 | 1.0956999999999999 | 1.1022000000000001 | 1.1399999999999999
 | 116862.4532 | 184484.8787 | 288893.8048 | 387512.3859 | 565898.239 | 754800.7336 | 967902.2664 | 1200342.9204 | 1541019.9494 | 1764539.5165 | 2021425.3837 | 2242996.846 | 2429469.8135 | 2667835.0522 | 2865123.0828 | 3076227.4734 | 3293039.9657 | 3464655.1314 | 3617940.0439 | 3767232.9981 | 3916326.9576 | 4012598.8465 | 4086418.0483 | 4139983.5763 | 4174446.6702 | 4193319.3194 | 4182003.7247 | 4154702.3088 | 4118358.9706 | 4088868.5021 | 4031000.6016 | 3971187.7509 | 3889991.0969 | 3818231.1469 | 3759495.5727 | 3671813.606 | 3590273.733 | 3497835.9158 | 3413024.1972 | 3356342.0697 | 3318296.7783 | 3289839.6705 | 3270568.9332 | 3232466.407 | 3213642.3981 | 3175592.0798 | 3147208.5919 | 3118440.6895 | 3099731.6156 | 3061973.3062 | 3024361.274 | 2977477.6483 | 2940272.3229 | 2912612.2154 | 2754440.3813
Chart: 
| 0.26519999999999999 | 0.28039999999999998 | 0.29780000000000001 | 0.31090000000000001 | 0.33040000000000003 | 0.3478 | 0.36520000000000002 | 0.3826 | 0.40649999999999997 | 0.42170000000000002 | 0.43909999999999999 | 0.45429999999999998 | 0.46739999999999998 | 0.48480000000000001 | 0.5 | 0.51739999999999997 | 0.53700000000000003 | 0.55430000000000001 | 0.57169999999999999 | 0.59130000000000005 | 0.61519999999999997 | 0.63480000000000003 | 0.65429999999999999 | 0.67390000000000005 | 0.69350000000000001 | 0.73040000000000005 | 0.75 | 0.77170000000000005 | 0.7913 | 0.80430000000000001 | 0.82609999999999995 | 0.84570000000000001 | 0.86960000000000004 | 0.8891 | 0.90429999999999999 | 0.92610000000000003 | 0.94569999999999999 | 0.96740000000000004 | 0.98699999999999999 | 1 | 1.0086999999999999 | 1.0152000000000001 | 1.0196000000000001 | 1.0283 | 1.0326 | 1.0412999999999999 | 1.0478000000000001 | 1.0544 | 1.0587 | 1.0673999999999999 | 1.0761000000000001 | 1.087 | 1.0956999999999999 | 1.1022000000000001 | 1.1399999999999999
 | -12957.2284 | -13970.7171 | -10253.8389 | -2015.7382 | 22573.1044 | 60477.4866 | 116495.3374 | 192767.1166 | 332642.1136 | 442910.0885 | 588721.6339 | 732088.3357 | 866543.771 | 1058903.7322 | 1237912.6198 | 1452894.3253 | 1704541.3167 | 1931971.169 | 2162803.769 | 2421082.1607 | 2726681.5716 | 2964190.6103 | 3184354.0687 | 3385791.6528 | 3564182.007 | 3828660.5322 | 3927698.505 | 4001792.2405 | 4035778.4424 | 4040880.5449 | 4018177.8537 | 3964574.5814 | 3857458.377 | 3736971.4958 | 3623078.9452 | 3430301.0121 | 3228611.3198 | 2975627.2307 | 2721824.8092 | 2540967.7732 | 2414606.4329 | 2317501.2471 | 2250489.6459 | 2115022.8493 | 2046644.4696 | 1905496.7718 | 1797657.1374 | 1686132.3983 | 1612400.14 | 1460704.2525 | 1305743.8738 | 1107180.7352 | 945318.5762 | 822510.1682 | 79180.1753
Wavelength
Wavelength
Wavelength
Detector Response
Blackbody Intensity
Net Measurement
Notes: As shown, non-linearities in the detector calibration are introduced because the measured signal is the net product of the detector response multiplied by the blackbody emission intensity. This can be compensated in the calibration.

## Slide 67: Wide-Bandwidth Detectors - IV
Using a wide-band detector can then require relating the integral of the blackbody function multiplied by the detector response across the bandwidth.
When measuring materials with emissivities that vary over wavelength, the data analysis becomes even more complex and requires the inclusion of the emissivity with wavelength.
So that the measured signal becomes a function of the intensity, emissivity, and the detector spectral response function:
And not just the intensity.
 |  | (31)
 |  | (31)
Notes: Using a wide-band detector can then require relating the integral of the blackbody function multiplied by the detector response across the bandwidth. When measuring materials with emissivities that vary over wavelength, the data analysis becomes even more complex and requires the inclusion of the emissivity with wavelength so that the measured signal becomes a function of the detector response, emissivity, and blackbody spectral intensity.
Using a wide-band detector can then require relating the integral of the blackbody function 𝑖_(𝑏,𝜆) (, 𝑇)multiplied by the detector response 𝐷()across the bandwidth. When measurements of materials with emissivities that vary over wavelength, the data analysis becomes even more complex and requires the inclusion of the emissivity 𝜀_ () with wavelength so that the measured signal becomes a function of the detector response, emissivity, and blackbody spectral intensity.
And not just 𝑖_(𝑏,𝜆) (, 𝑇).

## Slide 68: Wide-Bandwidth Detectors - V
Often, the analysis will use the average wavelength in the center of the band to calculate the blackbody intensity and assume a uniform intensity across the band, so that:
where:
 |  | (33)
 |  | (33)
 |  | (32)
 |  | (32)
Notes: Often, the analysis of wide-bandwidth detectors will use the average wavelength in the center of the band to calculate the blackbody intensity and assume a uniform intensity across the band. This can simplify the analysis but may lead to unaccounted for errors.

## Slide 69: Example 13 - I
E14: What is the non-linearity introduced by averaging the response of a detector covering the range of 1 to 4 m at the center wavelength over the range of temperatures from 1,500 to 4,000 K? Assume a linearly increasing detector spectral response function as shown below:
Chart: Wavelength, m
| 1 | 4 | 4
 | 0 | 1 | 0
Notes: Example 14. What is the non-linearity introduced by averaging the response of a detector covering the range of 1 to 4 m at the center wavelength over the range of temperatures from 1,500 to 4,000 K? Assume a linearly increasing detector spectral response function as shown.

## Slide 70: Example 13 - II
A14: The problem requires that we compare:
to:
where
and
So = 2.5 m and = 3 m. The comparison is shown on the following page.
Notes: The problem now requires that we compare the integral of spectral intensity multiplied by the detector spectral response function across the band to the product of the detector spectral response function and the spectral intensity at the average wavelength.

## Slide 71: Example 13-III
Note the non-linearity in the calibration and the high error at low intensities.
Chart: Actual Signal, W/m2-sr-m
| 30579.019904748624 | 38890.178072269191 | 48335.522382560972 | 58910.358738744297 | 70600.714977546741 | 83385.198384516843 | 97236.636018406105 | 112123.4971158635 | 128011.10617691663 | 144862.66149564477 | 162640.07751386153 | 181304.67101364932 | 200817.71135481371 | 221140.85416162648 | 242236.47643970198 | 264067.92934211408 | 286599.72290951887 | 309797.65522273746 | 333628.896617711 | 358062.0379737016 | 383067.11062085221 | 408615.58412917488 | 434680.34713244101 | 461235.67539508431 | 488257.19053226861 | 515721.81212557334
Signal | 40319.3822 | 51553.4142 | 64126.9875 | 77961.894 | 92975.7292 | 109085.6791 | 126211.0143 | 144274.6421 | 163203.9906 | 182931.4327 | 203394.4018 | 224535.3084 | 246301.3328 | 268644.1476 | 291519.6041 | 314887.4068 | 338710.7918 | 362956.2166 | 387593.0689 | 412593.3949 | 437931.6507 | 463584.474 | 489530.4775 | 515750.061 | 542225.2432
Error | -0.21688 | -0.13349 | -0.077323 | -0.039087 | -0.013068 | 0.0043976 | 0.015742 | 0.022628 | 0.026224 | 0.027366 | 0.026668 | 0.024582 | 0.021452 | 0.017536 | 0.013035 | 0.0081028 | 0.0028608 | -0.0025969 | -0.0081962 | -0.013879 | -0.019598 | -0.025318 | -0.031011 | -0.036653 | -0.042227 | -0.047721
Notes: The result is shown here. Again note the non-linearity in the calibration and the high error at low intensities.

## Slide 72: Wide-Bandwidth Detectors - V
When measurements are made of materials with emissivities that vary over wavelength, the data analysis becomes even more complex and requires the inclusion of the emissivity with wavelength:
 |  | (33)
 |  | (33)
Notes: When measurements are made of materials with emissivities that vary over wavelength, the data analysis becomes even more complex and requires the inclusion of the emissivity with wavelength as shown.

## Slide 73: Wide-Bandwidth Detectors - VI
When two or more detectors are used, the relationships for the unknown temperatures and spectral emissivities involve integrals over the wavelength band.
Similar to the narrow-band ratio method, we can devise an analogous way to determine temperature and a wavelength-averaged emissivity.
Notes: When the two or more detectors are used, the relationships for the unknown temperatures and spectral emissivities involve integrals over the wavelength band. Similar to the narrow-band ratio method, we can devise an analogous way to determine temperature and a wavelength-averaged emissivity.

## Slide 74: Wide-Bandwidth Detectors - VII
For two, wide-band detectors measuring equivalent blackbody temperatures of and , we can define an equivalent ratio method where we solve the following:
Where and are wavelength average emissivities. The next example will demonstrate this method.
 |  | (34)
 |  | (34)
Notes: For two, wide-band detectors measuring equivalent blackbody temperatures of and , we can define an equivalent ratio method where we solve the equation shown. Note however that this method is only an idealization if the emissivity is a strong function of wavelength. The symbols and are wavelength average emissivities. The next example will demonstrate this method.
For two, wide-band detectors measuring equivalent blackbody temperatures of 𝑇_1 and 𝑇_2, we can define an equivalent ratio method where we solve the equation shown. The symbols (ε_1 ) ̅ and (ε_2 ) ̅ are wavelength average emissivities. The next example will demonstrate this method.

## Slide 75: Wide-Bandwidth Detectors - VIII
Alternatively, we can approximate the wide-band signal using the average value across the band so that:
Given and we calculate that satisfies either Equation 34 or 35.
 |  | (35)
 |  | (35)
Notes: Alternatively, we can approximate the wide-band signal using the average value across the band as shown. Given and we calculate that satisfies either Equation 34 or 35.
Alternatively, we can approximate the wide-band signal using the average value across the band as shown. Given 𝑇_1 and 𝑇_2we calculate 𝑇 that satisfies either Equation 34 or 35.

## Slide 76: Example 14 - I
E14: A commercial Silicon/InGaAs sandwich detector has the detector response below. If the Silicon detector measures 3,013 K and the InGaAs measures 2,909 K equivalent blackbody temperatures, what are the average spectral emissivities for the two wavelengths and the true surface temperature assuming an emissivity ratio of 1.0?
Responsivity, A/W
0.6
0.5
0.4
0.3
0.2
0.1
0.0
Wavelength, m
0.35 0.45 0.55 0.65 0.75 0.85 0.95 1.05 1.15 1.25 1.35 1.45 1.55 1.65 1.75
Silicon
InGaAs
Notes: Example 14: A commercial Silicon/InGas sandwich detector has the detector response as shown. If the Silicon detector measures 3013 K and the InGas measures 2,909 K equivalent blackbody temperatures, what are the average spectral emissivities for the two wavelengths and the true surface temperature assuming an emissivity ratio of 1.0?

## Slide 77: Example 14 - II
Responsivity, A/W
0.6
0.5
0.4
0.3
0.2
0.1
0.0
Wavelength, m
0.35 0.45 0.55 0.65 0.75 0.85 0.95 1.05 1.15 1.25 1.35 1.45 1.55 1.65 1.75
Silicon
InGaAs
A14: We model the detector response function by the line segments below.
Notes: We model the detector response function by the line segments as shown.
We model the detector response function by the line segments as shown. The problem requires that we solve Equation 34 or 35 with (ε_𝑟 ) ̅=0.9 and 𝑇_1 = 3000 K and 𝑇_2 = 2800 K.

## Slide 78: Example 14 - III
The problem requires that we solve Equation 34 or 35 for the true surface temperature with
and = 3,013 K and = 2,909 K.
Integrated Wide Band Method
Approximate Averaged Band Method
 |  | (34)
 |  | (34)
 |  | (35)
 |  | (35)
Notes: The problem requires that we solve Equation 34 or 35 for the true surface temperature with an emissivity ratio and detector temperatures = 3,013 K and = 2,909 K.
The problem requires that we solve Equation 34 or 35 for the true surface temperature 𝑇 with an emissivity ratio (_𝑟 ) ̅=0.9 and detector temperatures 𝑇_1 = 3,000 K and 𝑇_2 = 2,950 K.

## Slide 79: Example 14 - IV
To ensure that the problem is self-consistent, we have calculated the equivalent blackbody temperatures from integration of the surface temperature across the wavelength band. The calculation uses the detector response functions as shown previously with an uniform emissivity over the wavelength range and equal to 0.7:
For , = and = .
 |  | (35)
 |  | (35)
Notes: To ensure that the problem is self-consistent, we calculate the equivalent blackbody temperatures from integration of the 3,200-K surface temperature across the wavelength band. The calculation uses the detector response functions as shown previously with an emissivity uniform over the wavelength range and equal to 0.7.
The calculated equivalent blackbody temperatures are 3,013 K and 2909 K for detectors 1 and 2, respectively,
To ensure that the problem is self-consistent, we calculate the equivalent blackbody temperatures 𝑇_ from integration of the surface temperature 𝑇 equal toy 3,200 K of across the wavelength band. The calculation uses the detector response functions as shown previously with an emissivity uniform over the wavelength range and equal to 0.7.
Using the function bb_tiibl, we can
To ensure that the probem is self-consistent, we calculate what the measured equivalent blackbody temperatures 𝑇_𝜆 are by solving the above equation given the detector response functions 𝐷(𝜆), the emissivity 𝜀 (independent of wavelength), and the spectral blackbody radiant intensity. The calculated blackbody temperatures for detectors 1 and 2 are 2,589 and 2,426 K, respectively.

## Slide 80: Example 14 - V
Parameter | Integrated Band | Averaged Band
, K | 3,200 | 3,134
 | 0.70 | 0.77
 | 0.70 | 0.76
Parameter | Integrated Band | Averaged Band
 | 3,200 | 3,134
 | 0.70 | 0.77
 | 0.70 | 0.76
Note that while the temperature is reasonably close between the two methods (66 K or 2%) the emissivities are less accurate, differing by about 10%.
= 1.0
Notes: The solution is shown here.

## Slide 81: Example 14 - VI
Parameter | Integrated Band | Averaged Band
, K | 3,200 | 3,134
 | 0.70 | 0.77
 | 0.70 | 0.76
Parameter | Integrated Band | Averaged Band
 | 3,200 | 3,134
 | 0.70 | 0.77
 | 0.70 | 0.76
Note that while the temperature is reasonably close between the two methods (66 K or 2%) the emissivities are less accurate, differing by about 10%.
= 1.0
Notes: The result for the integrated band method is exact because the problem was constructed to provide a surface temperature of exactly 3,200 K and an emissivity equal to 0.7 for this method.

## Slide 82: Example 14 - VII
Parameter | Integrated Band | Averaged Band
, K | 3,200 | 3,134
 | 0.70 | 0.77
 | 0.70 | 0.76
Parameter | Integrated Band | Averaged Band
 | 3,200 | 3,134
 | 0.70 | 0.77
 | 0.70 | 0.76
Note that while the temperature is reasonably close between the two methods (66 K or 2%) the emissivities are less accurate, differing by about 10%.
= 1.0
Notes: The temperature computed by the averaged band method is reasonably close between to the exact answer (within 66 K or 2%). However, the emissivities are not as accurate, differing by about 10%.

## Slide 83: Example 15 - I
E15: For the final problem in this section, we construct a problem that demonstrates not only the complexity in reducing data from
wide-band detectors, but also the difficulty in interpreting the data. Consider a dual-band instrument with spectral bands covering 1 to 4.5 m and 2 to 13.5 m. The detector response functions are shown.
Detector Response
Notes: Example 15: For the final problem, we construct a problem that demonstrates not only the complexity in reducing data from wide-band detectors, but also the difficulty in interpreting the data. Consider a dual-band instrument with spectral bands covering 1 to 4.5 m and 2 to 13.5 m. The detector response functions are shown.

## Slide 84: Example 15 - II
The material has a spectral emissivity as a function of wavelength as shown. The true surface temperature is 3,000 K.
Chart: Wavelength, m
| 0 | 14
 | 1 | 0.65
Notes: The material has a spectral emissivity as a decreasing function of wavelength as shown. The true surface temperature is 3,000 K. Note the varying emissivity generally cannot be accounted for in the analysis because seldom is it known exactly at the measured conditions.

## Slide 85: Example 15 - III
For the given true surface temperature , the calculated theoretically measured equivalent blackbody temperatures for each detector can be obtained by solving the following:
and for detectors 1 and 2 these are 2,941 and 2,841 K, respectively. If an emissivity ratio of 1 is assumed for the two detectors, what are the calculated inferred true surface temperature and spectral emissivities?
Notes: For the given true surface temperature , we can calculate the “theoretically measured” equivalent blackbody temperatures for each detector as shown. The calculated equivalent blackbody temperatures assuming emssiivities of 1 for detectors 1 and 2 are 2,949 and 2,719 K, respectively. If an emissivity ratio of 1 is assumed for the two detectors, what are the calculated true surface temperature and spectral emissivities?
For the given true surface temperature 𝑇, we can calculate the equivalent blackbody temperatures 𝑇_(𝜆_𝒊 )for each detector as shown. The calculated equivalent blackbody temperatures for detectors 1 and 2 are 2,949 and 2,719 K, respectively.

## Slide 86: Example 15 - IV
 | 1.0 |  | 1.16 | 
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
, K | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
 | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Actual Temperature = 3,000 K
Notes: A summary of the solutions is shown here. Note the wide variation in computed temperature and all results are far away from the true value of 3,000 K.

## Slide 87: Example 15 - V
 | 1.0 |  | 1.16 | 
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
, K | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
 | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Actual Temperature = 3,000 K
Notes: Using the wide-band, integrated method with Equation 34 assuming an emissivity ratio = 1, the calculated true surface temperature is 3,175 K with emissivities for the two detectors equal to 0.81 (routine bb_itratio). This is a 179-K error.
Using the wide-band, integrated method with Equation 34 assuming an emissivity ratio (ε_𝑟 ) ̅ = 1, the calculated true surface temperature is 3,175 K with emissivities for the two detectors equal to 0.81 (routine bb_itratio). This is a 179-K error.

## Slide 88: Example 15 - VI
 | 1.0 |  | 1.16 | 
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
, K | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
 | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Actual Temperature = 3,000 K
Notes: If an emissivity ratio of 1.16 is assumed (equal to the ratio of the emissivities at the mean detector wavelengths), then a surface temperature of 2,707 K is calculated and the error in surface temperature increases to 294 K. The corresponding emissivities are calculated to be 1.27 and 1.10 for detector 1 and 2, respectively. These are not physically realistic.

## Slide 89: Example 15 - VII
 | 1.0 |  | 1.16 | 
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
, K | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
 | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Actual Temperature = 3,000 K
Notes: Using the equivalent average wavelength assumption per Equation 35 with an emissivity ratio of 1 results in a calculated temperature of 3,127 K and emissivities for both detectors equal to 0.88.

## Slide 90: Example 15 - VIII
 | 1.0 |  | 1.16 | 
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
, K | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Parameter | Integrated Band | Averaged Band | Integrated Band | Averaged Band
 | 3,175 | 3,127 | 2,707 | 2,623
 | 0.81 | 0.88 | 1.27 | 1.29
 | 0.81 | 0.88 | 1.10 | 1.12
Actual Temperature = 3,000 K
Notes: When an emissivity ratio of 1.16 is used, the averaged band method gives a surface temperature of 2,623 K with emissivities for detector 1 and detector 2 being 1.29 and 1.12, respectively. Again, this s not physically realistic.. While the assumption of unity emissivity ratio comes closer to the correct surface temperature in both the wide-band and equivalent narrow-band methods, the result is fortuitous since the actual emissivity ratio across the two bands is not actually one.
The problem faced here with using wide-band detectors is that over such a wide range of wavelengths, it is difficult to choose the proper ratio since the wavelength dependence on spectral emissivity generally is not known.

## Slide 91: Rule 4
Unless other considerations dictate, avoid using wide-band detectors especially when accurate emissivities are needed.
Notes: Our final rule, Rule 4 is: Unless other considerations dictate, avoid using wide-band detectors especially when accurate emissivities are needed.

## Slide 92: Calibration
“Until you can measure something and express it in numbers, you have only the beginning of understanding.”
– Lord Kelvin
Notes: In this last section we will discuss calibration of radiation thermometers.

## Slide 93: Calibration
Detectors are calibrated against a source providing a known spectral intensity at a given wavelength.
However, it is often more convenient to express known calibration conditions in terms of an equivalent blackbody temperature, rather than a spectral intensity.
The calibrated spectral intensity and blackbody temperature are equivalent and can be used to determine the system response accounting for detector nonlinearities, transmission losses, and electrical gains.
Notes: Detectors are calibrated against a source providing a known spectral intensity at a given wavelength. However, it is often more convenient to express known calibration conditions in terms of an equivalent blackbody temperature, rather than a spectral intensity. The calibrated spectral intensity and blackbody temperature are equivalent and can be used to determine the system response accounting for detector nonlinearities, transmission losses, and electrical gains.

## Slide 94: Calibration Sources - I
Broadband Sources
  Blackbody Thermal Cavity – Near UV to IR
  Incandescent Lamp – Visible to Near IR
  Deuterium Arc Lamp – UV to Visible
Discrete Wavelength Sources
  Mercury Lamp
  Noble Gas (Neon, Argon, Krypton, Xenon) Lamp
Portable Blackbody Source
Deuterium Arc Lamp
Notes: There are several possible calibration sources that can used depending on the wavelength and temperature range required. These include broadband sources such as blackbody sources which cover the near UV to the IR, incandescent lamps used for the visible to near IR, and arc lamps, in the UV to visible bands. There are also discrete wavelength sources such as mercury and noble gas lamp sources that are calibrated for a single wavelength.

## Slide 95: Blackbody Source
A blackbody calibration source is an especially convenient and accurate calibration source.
Recall, a perfect blackbody is an ideal emitter and absorbs all incident radiation regardless of the spectral character of directionality of the incident radiation.
This behavior is simulated using a heated cavity designed to have an effective near unity emissivity.
The specified calibration conditions can be determined by measuring the temperature of the cavity, rather than the absolute spectral intensity, like other sources.
Notes: Note that the term “blackbody” source is not precisely the same as the definition of the blackbody used earlier in this lesson. But a blackbody calibration source is an especially convenient and accurate calibration source and produces radiation that closely approximates that of an ideal blackbody. Recall, a perfect blackbody is an ideal emitter and absorbs all incident radiation regardless of the spectral character of directionality of the incident radiation. This behavior is simulated using a heated cavity designed to have an effective near unity emissivity. The specified calibration conditions can be determined by measuring the temperature of the cavity, rather than the absolute spectral intensity, like other sources.

## Slide 96: Diagram of Blackbody Calibration Source
Notes: This figure shows a cutaway view of a typical blackbody calibration source. A well-insulated cavity made from a good emitter (although any material will theoretically work) and is heated, usually electrically. The temperature of the cavity is monitored with a contact temperature sensor (typically a thermocouple), and the emitted radiation from a narrow aperture provides a known intensity and spectral distribution of energy. Blackbody calibration sources are available from many commercial suppliers of optical equipment.

## Slide 97: Calibration Sources - II
Resistively Heated, High-temperature Blackbody Furnace – 1,000 to 3,500 K
Notes: Here we show pictures of an electrically heated blackbody capable of reaching temperatures approaching 3,500 K.

## Slide 98: Calibration Sources - III
Chart: Radiant Intensity, W/m2-m-sr
| 12.8 | 11.92 | 10.63 | 8.4600000000000009 | 7.95 | 6.38 | 5.8 | 4.32 | 3.2 | 4.26 | 5 | 3.83 | 0.53200000000000003
 | 32.3236 | 28.2026 | 24.5535 | 21.3289 | 18.4856 | 16.4587 | 13.7876 | 10.183 | 7.4434 | 10.9949 | 12.0449 | 8.8552 | 1.6041
Chart: Temperature, K
| 1273.1500000000001 | 1263.1500000000001 | 1253.1500000000001 | 1243.1500000000001 | 1233.1500000000001 | 1225.1500000000001 | 1213.1500000000001 | 1193.1500000000001 | 1173.1500000000001 | 1198.1500000000001 | 1204.1500000000001 | 1184.1500000000001 | 1084.1500000000001
 | 12.8 | 11.92 | 10.63 | 8.46 | 7.95 | 6.38 | 5.8 | 4.32 | 3.2 | 4.26 | 5 | 3.83 | 0.532
Blackbody calibration data for 0.656-m detector. The blackbody temperature was preset and the signal voltage measured. Plotting this versus the calculated radiant intensity produces a straight line of signal versus calculated intensity.
Notes: Detectors are calibrated to blackbody sources where the temperature of the cavity is measured. When the measured detector signal for a 0.656-micrometer detector is plotted versus temperature, the result is non-linear, as shown on the left figure. But if the measured spectral intensity is plotted versus the spectral intensity calculated from the Planck equations, the result can be more closely represented by a straight line.

## Slide 99: Calibration Uncertainty
All calibration sources have inherent error in the spectral intensity output.
For blackbody sources, this error manifests itself in the uncertainty in the cavity temperature measurement.
Uncertainty in the cavity intensity propagates error just like the emissivity.
Uncertainty in cavity temperature can be calculated from the quantity .
Notes: All calibration sources have inherent error in the spectral intensity output. For blackbody sources, this error manifests itself in the uncertainty in the cavity temperature measurement. Uncertainty in the cavity intensity propagates error just like the emissivity. Uncertainty in cavity temperature can be calculated from the quantity .
All calibration sources have inherent error in the spectral intensity output For blackbody sources, this error manifests itself in the uncertainty in the cavity temperature measurement. Uncertainty in the cavity intensity propagates error just like the emissivity. Uncertainty in cavity temperature can be calculated from the quantity 𝑑𝑇/𝑑𝜀.

## Slide 100: Example 16 - I
E16: A blackbody source is estimated to have an uncertainty of 10 K and is used to perform a single-point calibration on two detectors, 1) a short-wavelength, 0.5-m detector and 2) a longer-wavelength 3.0-m detector. These detectors then are used to make a measurement of an unknown material. The equivalent blackbody temperatures are measured to be 1,600 and 1,500 K, respectively. What is the value and the uncertainty in the derived emissivity at 3.0 m assuming the material has an estimated emissivity of 0.8 at 0.5 m?
Notes: Example 16. A blackbody source is estimated to have an uncertainty of 10 K and is used to perform a single-point calibration on two detectors, 1) a short-wavelength, 0.5-m detector and 2) a longer-wavelength 3 m detector. For an unknown material, the equivalent blackbody temperatures for each detector are measured to be 1,600 and 1,500 K, respectively. What is the value and the uncertainty in the derived emissivity at 3 m assuming the material has an emissivity of 0.8 at 0.5 m?

## Slide 101: Example 16 - II
A16: The true surface temperature from the measurement of the 0.5-m detector for = 1,600 K and = 0.8 is:
The calculated emissivity at 3.0 m is:
Assuming the temperature errors in the two calibration measurements are uncorrelated and random, the total uncertainty in the emissivity will be:
where
Notes: The true surface temperature temperature from the measurement of the 0.5-m detector for = 1,600 K and = 0.8 is calculated from Equation 13 and is 1,604 K. The calculated emissivity at 3 m from Equation 18 is 0.78. Assuming the temperature errors in the two calibration measurements are uncorrelated and random, the total uncertainty in the emissivity will be as shown where
The equivalent blackbody temperature from the measurement of the 0.5-m detector is for 𝑇_(𝜆_1 ) = 1,600 K and 𝜀_(𝜆_1 )= 0.8 is calculated from Equation 13 and is 1,604 K. The calculated emissivity at 3 m from Equation 18 is 0.78. Assuming the temperature errors in the two calibration measurements are uncorrelated and random, the total uncertainty in the emissivity will be as shown where where Δ𝑇_(𝜆_1 )=10K and 𝜀_(𝜆_1 )=0.1.

## Slide 102: Example 16 - III
or, in terms of sensitivities:
evaluating terms gives:
Notes: In terms of sensitivities, the relative error in emissivity is shown with the values of the individual terms.

## Slide 103: Example 16 - IV
So the estimated relative uncertainty is:
and therefore:
Detector 1 calibration
Detector 2 calibration
Notes: The final emissivity is found to be 0.78 plus minus 0.022.

## Slide 104: Example 16 - V
If the emissivity can only be estimated to 0.80.1, then error due to the uncertainty in the emissivity can be quantified using the calculated sensitivity of with . This is:
Therefore, the uncertainty in due to the uncertainty in is:
Note the uncertainty resulting from the calibration is on the same order as the uncertainty in the estimation of the emissivity at 0.5 m.
Notes: If the emissivity can only be estimated to 0.8 plus or minus 0.1, then error due to the uncertainty in the emissivity can be quantified using the calculated sensitivity of emissivity at wavelength 2 with respect the emissivity at wavelength 1. The estimated uncertainty is calculated to be 0.018. Note the uncertainty resulting from the calibration is on the same order as the estimation of the emissivity at 0.5 m.

## Slide 105: Closing
“If this fire determined by the sun, be received on the blackest known bodies, its heat will be long retain'd therein; and hence such bodies are the soonest and the strongest heated by the flame fire…”
– Hermann Boerhaave
A New Method of Chemistry, 2nd edition (1741), 262
Notes: Closing.

## Slide 106: Closing - I
Radiation thermometry is a useful technique for measuring the temperature of bodies that cannot be readily measured by contact sensors.
It is important to understand the uncertainty in radiation thermometry measurements and choose the characteristics of the instrumentation and the measurement method accordingly.
The most accurate temperature measurements are obtained when using short wavelength detectors.
The uncertainty in the derived spectral emissivity at a given wavelength decreases approximately inversely with the wavelength that is used, and therefore longer wavelength detectors provide more accurate emissivity measurements.
Notes: The following are the key points from this lesson:
Radiation thermometry is a useful technique for measuring the temperature of bodies that cannot be readily measured by contact sensors.
It is important to understand the uncertainty in radiation thermometry measurements and choose the characteristics of the instrumentation and the measurement method accordingly.
The most accurate temperature measurements are obtained when using short wavelength detectors.
The uncertainty in the derived spectral emissivity at a given wavelength decreases approximately inversely with the wavelength that is used, and therefore longer wavelength detectors provide more accurate emissivity measurements.

## Slide 107: Closing - II
Combining measurements of two or more detectors can allow the simultaneous determination of both the temperature and the emissivity.
However, the uncertainty in both the temperature and emissivity increase when the measurements of two detectors are combined.
Multi-spectral methods using a large number of detectors can often provide more accurate temperature and emissivity values over a wider range of wavelengths.
Prefer the use of narrow-band detectors over wide-band detectors unless other factors dictate.
Notes: Combining measurements of two or more detectors can allow the simultaneous determination of both the temperature and the emissivity.
However, the uncertainty in both the temperature and emissivity increase when the measurements of two detectors are combined.
Multi-spectral methods using a large number of detectors can often provide more accurate temperature and emissivity values over a wider range of wavelengths.
Prefer the use of narrow-band detectors over wide-band detectors unless other factors dictate.

## Slide 108: Summary of Radiative Thermometry Rules
Rule 1- When the emissivity is unknown and must be estimated, the most accurate surface temperature measurement is made when a detector with a wavelength as short as possible is used.
Rule 2 - The uncertainty in the spectral emissivity at a given wavelength decreases approximately inversely with the wavelength that is used.
Rule 3 - The optimal detector wavelengths for the ratio method are based on the two competing factors: 1) the need to keep the wavelength difference small to ensure an accurate assumed ratio, but 2) not too small that the effective wavelength increases and becomes too large.
Rule 4 - Unless other considerations dictate, avoid using wide-band detectors especially when accurate emissivities are needed.
Notes: To reiterate, there are four basic rules of radiative thermometry. Choosing the right detector type, wavelength, and measurement technique will ensure successful results.

## Slide 109: For Further Reading
Howell, J. R. , R. Siegel, and M. P. Mengüç, Thermal Radiation Heat Transfer, Fifth Edition, Taylor & Francis, New York, 2010.
DeWitt, D. P. and G. D. Nutter, Theory and Practice of Radiation Thermometry, John Wiley and Sons, New York, 1988.
Rogalski, A., “History of Infrared Detectors”, Opto-Electron. Rev., 20 no. 3, pp 279-308, 2012.
Cupa, R. and A. Rogalski, “Performance limitations of photon and thermal infrared detectors”, Opto-Electr. Rev., 5, no. 4, 1997.
DeWitt, D. P. and R. Rondeau, “Measurement of Surface Temperatures and Spectral Emissivities During Laser Irradiation”, J. Thermophysics, Vol. 3, No. 2, pp. 153-159, April 1989.
Johnson, P. E., D. P. DeWitt and R. E. Taylor, “Method for Measuring High Temperature Spectral Emissivity of Nonconducting Materials”, AIAA Journal, Vol. 19, No. 1, pp. 113-120, January 1981. Also AIAA Paper, 79-1037.
Water, W.R., J. H. Walker, A. T. Hattenberg, “NBS Measurement Services: Radiance Temperature Calibration, National Bureau of Standards Special Publication 250-7, Oct., 1987. Available at: http://www.nist.gov/calibrations/upload/sp250-7.pdf
Ready, J., “Fundamentals of Photonics, Module 1.6 Optical Detectors and Human Vision”, SPIE, http://spie.org/Documents/Publications/00%20STEP%20Module%2006.pdf.
Notes: The following references provide additional information on the topic of radiation thermometry and are recommended for those who wish to explore the subject further. Thank you for taking the time to explore this online NESC academy lesson.

## Slide 110: Acknowledgments
The author acknowledges funding for this work from the NASA NESC Passive Thermal Technical Committee. We are grateful for the comments and support provided by Mr. Steve Rickman, NESC Passive Thermal Technical Fellow and Mr. Daniel Ngyuen, Deputy of the Passive Thermal Technical Discipline Team. We also recognize the valuable comments and suggestions from Mr. Larry Hudson, Dr. William Ko, Mr. Chris Kostyk, and Mr. Matt Moholt of the Aerostructures Branch at NASA Armstrong Flight Research Center. Additional reviewers who provided many valuable comments also included Ms. Ruth Amundsen, Mr. Jentung Ku, Mr. David Gilmore, Mr. Richard Wear, and Mr. Duane Beach.
Notes: The author acknowledges funding for this work from the NASA NESC Passive Thermal Technical Committee. We are grateful for the comments and support provided by Mr. Steve Rickman, NESC Passive Thermal Technical Fellow and Mr. Daniel Ngyuen, Deputy of the Passive Thermal Technical Discipline Team. We also recognize the valuable comments and suggestions from Mr. Larry Hudson, Dr. William Ko, Mr. Chris Kostyk, and Mr. Matt Moholt of the Aerostructures Branch at NASA Armstrong Flight Research Center. Additional reviewers who provided many valuable comments also included Ms. Ruth Amundsen, Mr. Jentung Ku, Mr. David Gilmore, Mr. Richard Wear, and Mr. Duane Beach.

## Slide 111: FINE