## Slide 1
Aliasing Errors in Signal Analysis

## Slide 2
NASA Engineering & Safety Center (NESC) Academy
Dynamic Concepts, Inc. Huntsville, Alabama

## Slide 3
Tom Irvine
Email: tirvine@dynamic-concepts.com
Phone: (256) 922-9888

## Slide 4
Aliasing Introduction
Accelerometer data is digitized with a sample rate
The Nyquist frequency is one-half of the sample rate
A signal must not contain spectral components above the Nyquist frequency, otherwise aliasing will occur.
Aliasing is an error whereby high frequency energy is folded about the Nyquist frequency to lower frequencies
An analog lowpass filter may be used for anti-aliasing
An example of optical aliasing is shown in the next slide

## Slide 6
Consider a 200 Hz sine function.
The sine function is sampled at 2000 samples per second.
Thus there are 10 points per period.
The Nyquist frequency is 1000 Hz.
The Fourier magnitude is shown in the next slide.

## Slide 7
There is a distinct spectral line at 200 Hz as expected.

## Slide 8
Consider a 1800 Hz sine function.
The sine function is sampled at 2000 samples per second, as shown by the red markers. Thus there are 1.11 points per period.
The Nyquist frequency is 1000 Hz.
Aliasing occurs!

## Slide 9
The 1800 Hz signal appears as 200 Hz to the upcoming Fourier transform.

## Slide 10
The 1800 Hz signal is folded about the Nyquist frequency which is 1000 Hz.
The resulting energy is deposited at 200 Hz.

## Slide 11
Lowpass Filtering
Accelerometer
Signal Conditioner
with Analog
Lowpass Filter
Analog-to-Digital Converter
Aliasing can be prevented by lowpass filtering the analog data.
Consider a stage separation test or a launch vehicle flight.
The maximum expected frequency in the source energy is essentially unknown.
Thus, there is no proper means to set the sampling rate, other than setting it at some exceedingly high value.
The simple solution is to pass the analog data through a lowpass filter as shown in the flowchart.

## Slide 12
Butterworth Filter
The lowpass filter removes the high-frequency energy from the signal.
This filter is often called an "anti-aliasing" filter.
The filter can be part of the signal conditioning system.
Typically, a Butterworth filter is used.
The Butterworth filter has a roll-off which attenuates the signal by 3 dB at the cut-off frequency.
The cut-off frequency is typically set at, or slightly above, the maximum analysis frequency.

## Slide 13
Filter Guideline
The IES Handbook for Dynamic Data Acquisition and Analysis gives the following guidelines:
Let f c be the cutoff frequency
f N be the Nyquist frequency
A lowpass anti-aliasing filter with a cutoff rate of at least 60 dB/octave should be used prior to the analog-to-digital conversion of all dynamic data.
With a 60 dB/octave cutoff rate, the half-power point cutoff frequency of the filter should be set at fc < 0.6 fN
If the anti-aliasing filter has a more rapid cutoff rate, a higher cutoff frequency can be used, but the bound fc < 0.8 fN should never be exceeded.

## Slide 14
Sample Rate & Filter Examples
Sample Rate | Nyquist Freq | AA Filter Cut-off Freq | Typical Application
100 KHz | 50 KHz | 30 to 40 KHz | Mid or Far-field Pyroshock
5000 Hz | 2500 Hz | 1500 to 2000 Hz | Random Vibration
The following examples are mainly for launch vehicle flight telemetry.
The appropriate sample rate will vary on a case-by-case basis.

## Slide 15
aliasing.m
>> aliasing
aliasing.cpp ver 1.2 February 16, 2011
by Tom Irvine Email: tomirvine@aol.com
Select calculation type:
1=find resulting aliasing frequency
2=find candidate energy source frequencies
1
Enter the sampling rate ( samples/sec) 100
Enter the energy frequency (Hz) 70
aliasing frequency = 30 Hz

## Slide 16
>> aliasing
aliasing.cpp ver 1.2 February 16, 2011
by Tom Irvine Email: tomirvine@aol.com
Select calculation type:
1=find resulting aliasing frequency
2=find candidate energy source frequencies
2
Enter the sampling rate (samples/sec)
2000
Enter the resulting aliasing frequency (Hz)
200
Candidate energy frequencies (Hz)
1800
2200
3800
4200
5800
6200

## Slide 17
Waterfall FFT Launch Vehicle X Delta Velocity
Frequency (Hz)
Time (sec)
The sensor was from an Inertial Navigation System (INS). The data was sampled at 100 samples per second with no anti-aliasing filter. The waterfall FFT is given up to 50 Hz, the Nyquist frequency.
The spectral peaks from 25 to 35 Hz and from 50 to 60 seconds are due to aliasing about the Nyquist frequency. The source energy was a motor pressure oscillation that swept downward from 75 to 65 Hz.

## Slide 18
Simulated Aliasing, Example 1
SRS Q=10 | 
Natural Freq (Hz) | Peak Accel (G)
100 | 10
2000 | 1000
250K | 1000
The following examples shows that aliasing can cause up to 20 dB error in SRS plots, but that a massive amount of ultra-high-frequency energy is required for this to happen.
Consider a hypothetical shock pulse which satisfies the SRS in the following table.

## Slide 19
The top plot is a sample time history which reasonably satisfies the SRS specification for example 1. It is composed of damped sinusoids. The sample rate is 2.5 MHz
The bottom plot shows the same signal, decimated to a sample rate of 78.125 KHz.

## Slide 20
A close-up view is shown.

## Slide 21
The agreement is very good despite the aliasing present in the decimated data.

## Slide 22
Simulated Aliasing, Example 2
SRS Q=10 | 
Natural Freq (Hz) | Peak Accel (G)
100 | 10
2000 | 1000
250K | 50K
Repeat the previous example but with significantly higher amplitude at the last breakpoint.

## Slide 23
The top plot is a sample time history which reasonably satisfies the SRS specification for example 2. It is composed of damped sinusoids. The sample rate is 2.5 MHz
The bottom plot shows the same signal, decimated to a sample rate of 78.125 KHz.

## Slide 24
Aliasing occurs in the Decimated time history.

## Slide 25
The Decimated SRS is approximately 10 to 20 dB higher than the Original SRS.
The source of the error is aliasing.

## Slide 26
Program Summary
Reference
IEST-RP-DTE012.1 Handbook for Dynamic Data Acquisition and Analysis, Institute of Environmental Sciences, Illinois.
www.iest.org
Program
aliasing.m